## Work Summary Form
|  |  | Engineering World Health | [8 empty] | 111 South Highland, Suite 289
Memphis, TN 38111
901-624-2036
|  | Complete electronically if possible | [9 empty] | www.ewh.org
* REQUIRED FIELD
| *DATE |  |  |  | *ENGINEER NAME(S) |  |  |  |  | Siqi Ngan, David Houck
| *COUNTRY | Nicaragua |  |  | *Hospital/CITY |  |  |  |  | Velez Paiz, Managua
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION | [7 empty] | Results
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
GUIDE
1 | Infant Warmer (radiant or other) | Air Shields | PM78-1 | HX13387 |  |  | x |  |  |  |  | Resistor had lost connection | x
2 | Electrosurgery Unit (ESU) | n/a | n/a | n/a |  |  | x |  |  |  |  | Built an adaptor for an electrode for a different plug | x
3 | Other | Mada Medical Products | DOL-702A-AA | NN 013309 | x | [6 empty] | Cleaned inside of the nebulizer | x
4 | other | Mada Medical Products | DOL-702A-AA | LR37209 | x | [6 empty] | Cleaned inside of the nebulizer | x
5 | Infant Warmer (radiant or other) | Nakamura Medical IND. Co. LTD | NIW-2000 | S19301241 |  |  | x |  |  |  |  | Wasn't heating up: cleaned fuses, replaced transistor but still only functions from controls 6-8 (1-5 = 0), wrote a note to explain machine functions | x
6 | Electrosurgery Unit (ESU) | Valleylab | Force 4B | POF 2208B | [6 empty] | x |  | x
7 | Infant Warmer (radiant or other) | Airshields | PM78-1 | CPO 3007 |  |  |  | x |  |  |  | Missing a bottom wheel: Replaced the bottom section of the warmer with another from the bodega | x
8 | Infusion pumps | Becton Dickinson | Program1 | 15151343 |  |  | x |  |  |  |  | Contacts on the front panel needed to be cleaned | x
9 | Infant Warmer (radiant or other) | Nakamura Medical IND. Co. LTD | NIW-2000 | 9301244 |  |  | x |  |  |  |  | Resistor had lost connection | x
10 | Infant Warmer (radiant or other) | Airshields | PM78-1 | YG17496 |  |  |  | x |  |  |  | Switch was lubricated for better contact | x
11 | Incubator (infant) | Infant | YP-90 | 41211970 |  |  |  | x |  |  |  | Door latch had been removed, attached two latches on either side, cleaned machine | x
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96
97
98
99
100

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Photocopier
Phototherapy device
Printer
Projector
Pulse Oximeter
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Other

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
Velez Paiz, Managua01 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 1 | Infant Warmer (radiant or other) | Air Shields | PM78-1 | HX13387 | 0 | 0 | x | 0 | 0 | 0 | 0 | Resistor had lost connection | x | 0
Velez Paiz, Managua02 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 2 | Electrosurgery Unit (ESU) | n/a | n/a | n/a | 0 | 0 | x | 0 | 0 | 0 | 0 | Built an adaptor for an electrode for a different plug | x | 0
Velez Paiz, Managua03 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 3 | Other | Mada Medical Products | DOL-702A-AA | NN 013309 | x | 0 | 0 | 0 | 0 | 0 | 0 | Cleaned inside of the nebulizer | x | 0
Velez Paiz, Managua04 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 4 | other | Mada Medical Products | DOL-702A-AA | LR37209 | x | 0 | 0 | 0 | 0 | 0 | 0 | Cleaned inside of the nebulizer | x | 0
Velez Paiz, Managua05 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 5 | Infant Warmer (radiant or other) | Nakamura Medical IND. Co. LTD | NIW-2000 | S19301241 | 0 | 0 | x | 0 | 0 | 0 | 0 | Wasn't heating up: cleaned fuses, replaced transistor but still only functions from controls 6-8 (1-5 = 0), wrote a note to explain machine functions | x | 0
Velez Paiz, Managua06 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 6 | Electrosurgery Unit (ESU) | Valleylab | Force 4B | POF 2208B | 0 | 0 | 0 | 0 | 0 | 0 | x | 0 | x | 0
Velez Paiz, Managua07 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 7 | Infant Warmer (radiant or other) | Airshields | PM78-1 | CPO 3007 | 0 | 0 | 0 | x | 0 | 0 | 0 | Missing a bottom wheel: Replaced the bottom section of the warmer with another from the bodega | x | 0
Velez Paiz, Managua08 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 8 | Infusion pumps | Becton Dickinson | Program1 | 15151343 | 0 | 0 | x | 0 | 0 | 0 | 0 | Contacts on the front panel needed to be cleaned | x | 0
Velez Paiz, Managua09 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 9 | Infant Warmer (radiant or other) | Nakamura Medical IND. Co. LTD | NIW-2000 | 9301244 | 0 | 0 | x | 0 | 0 | 0 | 0 | Resistor had lost connection | x | 0
Velez Paiz, Managua010 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 10 | Infant Warmer (radiant or other) | Airshields | PM78-1 | YG17496 | 0 | 0 | 0 | x | 0 | 0 | 0 | Switch was lubricated for better contact | x | 0
Velez Paiz, Managua011 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 11 | Incubator (infant) | Infant | YP-90 | 41211970 | 0 | 0 | 0 | x | 0 | 0 | 0 | Door latch had been removed, attached two latches on either side, cleaned machine | x | 0
Velez Paiz, Managua012 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua013 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua014 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 14 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua015 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 15 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua016 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 16 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua017 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua018 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua019 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua020 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua021 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 21 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua022 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua023 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 23 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua024 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua025 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua026 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 26 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua027 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 27 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua028 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 28 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua029 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 29 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua030 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua031 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 31 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua032 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 32 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua033 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua034 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua035 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua036 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua037 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua038 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua039 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua040 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua041 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua042 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua043 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua044 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua045 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua046 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua047 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua048 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua049 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 49 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua050 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua051 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua052 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 52 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua053 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua054 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 54 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua055 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua056 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua057 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 57 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua058 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 58 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua059 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 59 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua060 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua061 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 61 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua062 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 62 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua063 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua064 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 64 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua065 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua066 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 66 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua067 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua068 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua069 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 69 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua070 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 70 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua071 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 71 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua072 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 72 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua073 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 73 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua074 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua075 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua076 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 76 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua077 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua078 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 78 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua079 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 79 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua080 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua081 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 81 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua082 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 82 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua083 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 83 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua084 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 84 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua085 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 85 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua086 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 86 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua087 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 87 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua088 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 88 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua089 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua090 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua091 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua092 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua093 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 93 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua094 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 94 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua095 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua096 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 96 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua097 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 97 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua098 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 98 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua099 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Velez Paiz, Managua0100 | Velez Paiz, Managua | 0 | Siqi Ngan, David Houck | Nicaragua | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0