## Sheet1
|  |  | Engineering World Health | [8 empty] | 111 South Highland, Suite 289
Memphis, TN 38111
901-624-2036
|  | Complete electronically if possible | [9 empty] | www.ewh.org
* REQUIRED FIELD
| *DATE | 2012-08-06 |  |  | *ENGINEER NAME(S) |  |  |  |  | Morganne Gagne, Grace Cartner
| *COUNTRY | Tanzania |  |  | *Hospital/CITY |  |  |  |  | Mt. Meru Hospital, Arusha
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION | [7 empty] | Results
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric | Mechanical | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
1 | oxygen concentrator | Nuvo | Lite | [7 empty] | X | difficult to work with when opened |  | x
2 | suction machine | Fazzini |  |  |  | x | [5 empty] | replaced fuse, rewired motor | X
3 | oxygen concentrator | WD6F |  |  |  |  | X |  |  |  |  | replaced plug | X
4 | oxygen concentrator | New Life | Elite | [6 empty] | x |  | worked when plugged in | X
5 | other: power strip | [5 empty] | x |  |  |  |  | rewired | X
6 | wheelchair | [6 empty] | X |  |  |  | from TB ward, used electrical tape to make breaks, still wanted rubber for wheels | X
7 | other: drip stand | [6 empty] | x |  |  |  | used steel epoxy, still wobbles, Nuru said to repair more important things | X
8 | other: IR lamp | Microtizer | MT-150 |  |  |  | x |  |  |  |  | missing fuse, replaced plug | X
9 | nebulizer | Omron | NE-C28 | [6 empty] | x |  | working | X
10 | nebulizer | Omron | NE-C29 | [7 empty] | x | missing parts |  | X
11 | other: chair | [6 empty] | x |  |  |  | base was disconnected - put back together | X
12 | bili lights | PhotoTherapie | 800 |  |  |  | x |  |  |  |  | replaced old German plug with African plug | X
13 | centrifuge |  | DSC-300SD |  |  | x | [5 empty] | power is going to motor, motor isn't running, replaced carbon brushes |  | x
14 | blood/glucose testers | Optimum | Xceed |  |  |  | x |  |  |  |  | 2/6 needed batteries, 4/6 beyond repair | X
15 | oxygen concentrator | [8 empty] | x |  | brand new, from post-natal ward, needed adapter |  | X
16 | spygomanometer | [6 empty] | x |  |  |  | hole in tubing |  | X
17 | EBI | BioMET | [7 empty] | x |  | beyond repair, exceeded lifetime of device |  | x
18 | pipets | Huawei, HTL, Volac | [5 empty] | x |  |  |  | mechanically damaged and need calibration, beyond our scope of repair |  | X
19 | adapter | Volex |  |  |  |  | x |  |  |  |  | replaced fuse | X
20 | centrifuge | Baxter | Stat-60 |  |  |  | x |  |  |  |  | replaced brushes on motor, fixed switch, no power getting to motor, think problem is with interlocking system |  | X
21 | voltage stabilizer | [5 empty] | X |  |  |  |  | missing capacitors, would need to find similar stabilizer to find what type. Decided not to fix because it was a personal stabilizer |  | x
22 | suction machine | Con-Trex | SM-180 | 1095B01004 |  |  | X |  |  |  |  | put in new switch and rewired | x
23 | kettle | [5 empty] | X |  |  |  |  | wired around broken temperature sensor |  | x