## Work Summary Form
Engineering World Health
www.ewh.org
Complete electronically if possible
* REQUIRED FIELD
| *DATE |  |  |  | *ENGINEER NAME(S)
| *COUNTRY |  |  |  | *Hospital/CITY
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION | [7 empty] | Results
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
GUIDE
EXAMPLE1 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-5000 | 00W295 | [5 empty] | 1 |  | Machine was decommissioned previously |  | 1
EXAMPLE2 | Centrifuge | Gemmy | KHT-430 | 313139 |  |  |  | 1 |  |  |  | Rubber Seal was stiff, cut down to ensure fit | 1
1 | BP Cuff | Riester | [5 empty] | 1 |  |  |  | Leak in gauge and the needle sticks | 1
2 | BP Cuff | Riester | [8 empty] | 1 | miscalibrated but working | 1
3 | BP Cuff | [6 empty] | 1 |  |  |  | valve lets out air too quickly, cuff messed up | 1
4 | BP Cuff | Riester | [5 empty] | 1 |  |  |  | pump is hard to pump | 1
5 | BP Cuff | Riester | [5 empty] | 1 |  |  |  | clogged filter | 1
6 | BP Cuff | F.Bosch | 123 | [7 empty] | 1 | calibration error | 1
7 | BP Cuff | Riester | [5 empty] | 1 |  |  |  | clogged filter | 1
8 | BP Cuff | Riester | [5 empty] | 1 |  |  |  | knob comes off | 1
9 | BP Cuff | Riester | [8 empty] | 1 | broken |  | 1
10 | BP Cuff | [9 empty] | 1 | zero is at 280, miscalibrated | 1
11 | Dental Drilling Machine | Quayle |  | 138097-1497 |  |  |  | 1 |  |  |  | missing tubing connecting to compressor, needed maintenance (removing water and cleaning), fixed the water tank, attached the holders for the drilling heads | 1
12 | Autoclave | Fullers Medical | 1516 | [6 empty] | 1 |  | actually this was working and they didn’t realize, we just tested it | 1
13 | Autoclave | All American | 75X | 3231 |  |  |  | 1 |  |  |  | missing a rubber seal | 1
14 | Microscope | Olympus | CH30 |  |  |  | 1 |  |  |  |  | blown bulb, bought a new one | 1
15 | Fridge | Electrolux | TCW 1151 | [5 empty] | 1 |  |  | we dotn know whats wrong with this one |  | 1
16 | Mixer | Vivadent |  |  |  | 1 | [5 empty] | Burned coils |  | 1
17 | Mixer |  |  |  |  | 1 | [5 empty] | Burned coils |  | 1
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96
97
98
99
100

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Photocopier
Phototherapy device
Printer
Projector
Pulse Oximeter
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Other

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
Hospital Amistad Japon, GRANADA29/07/121 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 1 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-5000 | 00W295 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Machine was decommissioned previously | 0 | 1
Hospital Amistad Japon, GRANADA29/07/122 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 2 | Centrifuge | Gemmy | KHT-430 | 313139 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Rubber Seal was stiff, cut down to ensure fit | 1 | 0
Hospital Amistad Japon, GRANADA29/07/123 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 3 | Ultrasound Machine (imaging) | Diagnostic Ultrasound System | 4900 606-LV300 | BCA1146 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Machine was decommissioned previously | 0 | 1
Hospital Amistad Japon, GRANADA29/07/124 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 4 | Ultrasound Machine (imaging) | GE Medical Systems | 2205675 | 65361SM8 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Machine was decommissioned previously | 0 | 1
Hospital Amistad Japon, GRANADA29/07/125 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 5 | Autoclave (lab, surgery, and other) | HillMed | HM-Armada 1045 | 301575 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Repaired door seals | 1 | 0
Hospital Amistad Japon, GRANADA29/07/126 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 6 | Autoclave (lab, surgery, and other) | Napco | 8000 DSE | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Cleaned electrical connections | 1 | 0
Hospital Amistad Japon, GRANADA29/07/127 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 7 | Dental Drilling Machine | Actus | TU-85N | 6647 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Air and Water tubing was clogged | 1 | 0
Hospital Amistad Japon, GRANADA29/07/128 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 8 | Dental Drilling Machine | Actus | TU-85N | 6646 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Foot pedal was previously unresponsive | 1 | 0
Hospital Amistad Japon, GRANADA29/07/129 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 9 | Centrifuge | Gemmy | KHT-M30 | 313471 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Carbon Brushes were replaced | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1210 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 10 | Respiration monitor (apnea monitor) | Pulma-Aide | 0 | D3C4551 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Air tubing was clogged | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1211 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 11 | Respiration monitor (apnea monitor) | Sunrise | 0 | 14-6-4-315-4 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Air tubing and filter were clogged | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1212 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 12 | Aspirator/Suction Machine | Gomco | Premier Series | 1240 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Preventative maintenance on motor | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1213 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 13 | Respiration monitor (apnea monitor) | Memmert | 0 | WB7 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Air tubing was clogged | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1214 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 14 | Respiration monitor (apnea monitor) | Sunrise | 5650D | D35973161 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Preventative maintenance | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1215 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 15 | Respiration monitor (apnea monitor) | GAST | 0 | 1206605309 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Preventative maintenance | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1216 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 16 | Aspirator/Suction Machine | Nellcor Puriton Bennett | Radiation Adult Star 1500 | 4200071485 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Ran self-diagnostic test, Preventative mainenance | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1217 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 17 | Respiration monitor (apnea monitor) | GAST | NO180 | 1105625805 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Cleaned electrical connections | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1218 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 18 | Aspirator/Suction Machine | TAWANISHI | Tipo-1000 | 6313 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | Replace tubing, reconnect electrical connections, cleaned system | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1219 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 19 | Distiller | Kayagaki | WS-5 | 49865 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Balanced to make less noisy | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1220 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 20 | Centrifuge | Gemmy | PLC-05 | 702949 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Reconnect wires | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1221 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 21 | X-Ray Machine | Siemens | SC3-3 | 8C34-30 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Preventative maintenance | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1222 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 22 | X-Ray Machine | Siemens | SC3-3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1223 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 23 | Aspirator/Suction Machine | Gamco | 6036 | P-7123 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Increased suction, fixed wheels on bottom | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1224 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 24 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1225 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 25 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1226 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 26 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1227 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 27 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1228 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 28 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Preventative maintenance | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1229 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 29 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Replaced | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1230 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 30 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Replaced | 0 | 1
Hospital Amistad Japon, GRANADA29/07/1231 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 31 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Cleaned fan components | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1232 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 32 | Air Conditioner | Carrier | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | Cleaned fan components | 1 | 0
Hospital Amistad Japon, GRANADA29/07/1233 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1234 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1235 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1236 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1237 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1238 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1239 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1240 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1241 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1242 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1243 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1244 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1245 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1246 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1247 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1248 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1249 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 49 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1250 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1251 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1252 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 52 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1253 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1254 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 54 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1255 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1256 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1257 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 57 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1258 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 58 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1259 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 59 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1260 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1261 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 61 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1262 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 62 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1263 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1264 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 64 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1265 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1266 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 66 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1267 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1268 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1269 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 69 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1270 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 70 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1271 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 71 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1272 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 72 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1273 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 73 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1274 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1275 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1276 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 76 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1277 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1278 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 78 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1279 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 79 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1280 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1281 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 81 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1282 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 82 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1283 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 83 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1284 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 84 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1285 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 85 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1286 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 86 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1287 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 87 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1288 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 88 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1289 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1290 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1291 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1292 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1293 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 93 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1294 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 94 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1295 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1296 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 96 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1297 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 97 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1298 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 98 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/1299 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital Amistad Japon, GRANADA29/07/12100 | Hospital Amistad Japon, GRANADA | 29/07/12 | Scott Tyler, Anish Raman | Nicaragua | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0