## Work Summary Form
|  |  | Engineering World Health | [8 empty] | 111 South Highland, Suite 289
Memphis, TN 38111
901-624-2036
|  | Complete electronically if possible | [9 empty] | www.ewh.org
* REQUIRED FIELD
| *DATE |  |  |  | *ENGINEER NAME(S)
| *COUNTRY |  |  |  | *Hospital/CITY
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION | [7 empty] | Results
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
GUIDE
1 | UPS (battery backup for computer) | Syanex | VGD-700 | 11484 | [6 empty] | 1 | Gave a Bus fault error on the LCD screen , the manual said it was a problem with the internal DC voltage Bus |  | 1
2 | Lamp, examination | Titanox | SRL |  |  |  |  | 1 |  |  |  | mechanical lamp head broke the electrical connection of the lamp | 1
3 | X-Ray Machine | Siemens | OPTIPHOS 135/30/55R | 2274 |  |  | 1 |  |  |  |  | changed the light bulb in the machine | 1
4 | Autoclave (lab, surgery, and other) | Chinese | YXQ.WY21.600 | RZZP-065-2006 |  |  |  |  | 1 |  |  | connected the wire directly to the outlet | 1
5 | Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable) | _ | [5 empty] | 1 |  |  |  | needle not correctly alligned | 1
6 | Incubator (infant) | Chinese | [6 empty] | 1 |  |  | control panel could not be made to receive power |  | 1
7 | Other | (Stabilizer) | SLC-1200FE-8135 | :09020110 |  |  | 1 |  |  |  |  | wiring for the output was loose | 1
8 | Other | (Stabilizer) | SLC-1200FE-8135 | :09020115 |  |  | 1 |  |  |  |  | wiring for the output was loose | 1
9 | Hot Plate (laboratory, but not kitchen) | _ |  |  |  |  | 1 |  |  |  |  | changed the plug and reconnected the wire to the stove. | 1
10 | Aspirator/Suction Machine | Polivac | B4/SL7 | 0409-0823 |  |  | 1 |  |  |  |  | bad wiring and tube leaks | 1
11 | Aspirator/Suction Machine | Chinese | YB-LX-3 | 047054810071RZM | [6 empty] | 1 | Relay to allow suction was not working. We had to bypass it. | 1
12 | Infant Warmer (radiant or other) | _ | HKN-934 | 25080302010 | [6 empty] | 1 | Error E02 shown on LCD screen could not find manual to discern what that meant. Took the heating lamp off and used the base with other warmers. | 1
13 | Distiller | [6 empty] | 1 |  |  |  | bad pressure switch had to bypass it. | 1
14 | Lamp, examination | [5 empty] | 1 |  |  |  |  | burned out bulb and bad wiring. | 1
15 | Lamp, examination | [5 empty] | 1 |  |  |  |  | burned out bulb and bad wiring. | 1
16 | Aspirator/Suction Machine | EuroDel | [8 empty] | 1 | burned out diodes on the circuit board changed the burned out diodes but still no suction looked at the motor but could not figure out the problem. |  | 1
17 | Other | Roll Mixer | [8 empty] | 1 | The knob on the mixer does not reduce the speed of the roller when it is supposed to. We think the potentiometer might be at fault because it is part of what controls the speed of the mixer but there is nothing to suggest it is. |  | 1
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96
97
98
99
100

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Photocopier
Phototherapy device
Printer
Projector
Pulse Oximeter
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Other

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
001 | 0 | 0 | 0 | 0 | 1 | UPS (battery backup for computer) | Syanex | VGD-700 | 11484 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Gave a Bus fault error on the LCD screen , the manual said it was a problem with the internal DC voltage Bus | 0 | 1
002 | 0 | 0 | 0 | 0 | 2 | Lamp, examination | Titanox | SRL | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | mechanical lamp head broke the electrical connection of the lamp | 1 | 0
003 | 0 | 0 | 0 | 0 | 3 | X-Ray Machine | Siemens | OPTIPHOS 135/30/55R | 2274 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | changed the light bulb in the machine | 1 | 0
004 | 0 | 0 | 0 | 0 | 4 | Autoclave (lab, surgery, and other) | Chinese | YXQ.WY21.600 | RZZP-065-2006 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | connected the wire directly to the outlet | 1 | 0
005 | 0 | 0 | 0 | 0 | 5 | Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable) | _ | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | needle not correctly alligned | 1 | 0
006 | 0 | 0 | 0 | 0 | 6 | Incubator (infant) | Chinese | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | control panel could not be made to receive power | 0 | 1
007 | 0 | 0 | 0 | 0 | 7 | Other | (Stabilizer) | SLC-1200FE-8135 | :09020110 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | wiring for the output was loose | 1 | 0
008 | 0 | 0 | 0 | 0 | 8 | Other | (Stabilizer) | SLC-1200FE-8135 | :09020115 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | wiring for the output was loose | 1 | 0
009 | 0 | 0 | 0 | 0 | 9 | Hot Plate (laboratory, but not kitchen) | _ | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | changed the plug and reconnected the wire to the stove. | 1 | 0
0010 | 0 | 0 | 0 | 0 | 10 | Aspirator/Suction Machine | Polivac | B4/SL7 | 0409-0823 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | bad wiring and tube leaks | 1 | 0
0011 | 0 | 0 | 0 | 0 | 11 | Aspirator/Suction Machine | Chinese | YB-LX-3 | 047054810071RZM | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Relay to allow suction was not working. We had to bypass it. | 1 | 0
0012 | 0 | 0 | 0 | 0 | 12 | Infant Warmer (radiant or other) | _ | HKN-934 | 25080302010 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Error E02 shown on LCD screen could not find manual to discern what that meant. Took the heating lamp off and used the base with other warmers. | 1 | 0
0013 | 0 | 0 | 0 | 0 | 13 | Distiller | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | bad pressure switch had to bypass it. | 1 | 0
0014 | 0 | 0 | 0 | 0 | 14 | Lamp, examination | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | burned out bulb and bad wiring. | 1 | 0
0015 | 0 | 0 | 0 | 0 | 15 | Lamp, examination | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | burned out bulb and bad wiring. | 1 | 0
0016 | 0 | 0 | 0 | 0 | 16 | Aspirator/Suction Machine | EuroDel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | burned out diodes on the circuit board changed the burned out diodes but still no suction looked at the motor but could not figure out the problem. | 0 | 1
0017 | 0 | 0 | 0 | 0 | 17 | Other | Roll Mixer | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | The knob on the mixer does not reduce the speed of the roller when it is supposed to. We think the potentiometer might be at fault because it is part of what controls the speed of the mixer but there is nothing to suggest it is. | 0 | 1
0018 | 0 | 0 | 0 | 0 | 18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0019 | 0 | 0 | 0 | 0 | 19 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0020 | 0 | 0 | 0 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0021 | 0 | 0 | 0 | 0 | 21 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0022 | 0 | 0 | 0 | 0 | 22 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0023 | 0 | 0 | 0 | 0 | 23 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0024 | 0 | 0 | 0 | 0 | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0025 | 0 | 0 | 0 | 0 | 25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0026 | 0 | 0 | 0 | 0 | 26 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0027 | 0 | 0 | 0 | 0 | 27 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0028 | 0 | 0 | 0 | 0 | 28 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0029 | 0 | 0 | 0 | 0 | 29 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0030 | 0 | 0 | 0 | 0 | 30 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0031 | 0 | 0 | 0 | 0 | 31 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0032 | 0 | 0 | 0 | 0 | 32 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0033 | 0 | 0 | 0 | 0 | 33 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0034 | 0 | 0 | 0 | 0 | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0035 | 0 | 0 | 0 | 0 | 35 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0036 | 0 | 0 | 0 | 0 | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0037 | 0 | 0 | 0 | 0 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0038 | 0 | 0 | 0 | 0 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0039 | 0 | 0 | 0 | 0 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0040 | 0 | 0 | 0 | 0 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0041 | 0 | 0 | 0 | 0 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0042 | 0 | 0 | 0 | 0 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0043 | 0 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0044 | 0 | 0 | 0 | 0 | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0045 | 0 | 0 | 0 | 0 | 45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0046 | 0 | 0 | 0 | 0 | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0047 | 0 | 0 | 0 | 0 | 47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0048 | 0 | 0 | 0 | 0 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0049 | 0 | 0 | 0 | 0 | 49 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0050 | 0 | 0 | 0 | 0 | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0051 | 0 | 0 | 0 | 0 | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0052 | 0 | 0 | 0 | 0 | 52 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0053 | 0 | 0 | 0 | 0 | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0054 | 0 | 0 | 0 | 0 | 54 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0055 | 0 | 0 | 0 | 0 | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0056 | 0 | 0 | 0 | 0 | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0057 | 0 | 0 | 0 | 0 | 57 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0058 | 0 | 0 | 0 | 0 | 58 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0059 | 0 | 0 | 0 | 0 | 59 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0060 | 0 | 0 | 0 | 0 | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0061 | 0 | 0 | 0 | 0 | 61 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0062 | 0 | 0 | 0 | 0 | 62 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0063 | 0 | 0 | 0 | 0 | 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0064 | 0 | 0 | 0 | 0 | 64 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0065 | 0 | 0 | 0 | 0 | 65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0066 | 0 | 0 | 0 | 0 | 66 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0067 | 0 | 0 | 0 | 0 | 67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0068 | 0 | 0 | 0 | 0 | 68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0069 | 0 | 0 | 0 | 0 | 69 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0070 | 0 | 0 | 0 | 0 | 70 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0071 | 0 | 0 | 0 | 0 | 71 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0072 | 0 | 0 | 0 | 0 | 72 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0073 | 0 | 0 | 0 | 0 | 73 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0074 | 0 | 0 | 0 | 0 | 74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0075 | 0 | 0 | 0 | 0 | 75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0076 | 0 | 0 | 0 | 0 | 76 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0077 | 0 | 0 | 0 | 0 | 77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0078 | 0 | 0 | 0 | 0 | 78 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0079 | 0 | 0 | 0 | 0 | 79 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0080 | 0 | 0 | 0 | 0 | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0081 | 0 | 0 | 0 | 0 | 81 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0082 | 0 | 0 | 0 | 0 | 82 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0083 | 0 | 0 | 0 | 0 | 83 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0084 | 0 | 0 | 0 | 0 | 84 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0085 | 0 | 0 | 0 | 0 | 85 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0086 | 0 | 0 | 0 | 0 | 86 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0087 | 0 | 0 | 0 | 0 | 87 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0088 | 0 | 0 | 0 | 0 | 88 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0089 | 0 | 0 | 0 | 0 | 89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0090 | 0 | 0 | 0 | 0 | 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0091 | 0 | 0 | 0 | 0 | 91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0092 | 0 | 0 | 0 | 0 | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0093 | 0 | 0 | 0 | 0 | 93 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0094 | 0 | 0 | 0 | 0 | 94 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0095 | 0 | 0 | 0 | 0 | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0096 | 0 | 0 | 0 | 0 | 96 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0097 | 0 | 0 | 0 | 0 | 97 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0098 | 0 | 0 | 0 | 0 | 98 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
0099 | 0 | 0 | 0 | 0 | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
00100 | 0 | 0 | 0 | 0 | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0