## Work Summary Form
|  | Engineering World Health Work Summary | [9 empty] | The Prizery, Suite 200
302 East Pettigrew Street
Durham, NC 27701
|  | Complete electronically if possible | [9 empty] | www.ewh.org
* REQUIRED FIELD
| *DATE | 14. July 2015 |  |  | *ENGINEER NAME(S) |  |  |  |  | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen
| *COUNTRY | Tanzania |  |  | *Hospital/CITY |  |  |  |  | KCMC Moshi |  |  | 65 | 18 | 0.78313
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION | [7 empty] | Results |  | Details
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned | Location | Year of Manufacture
GUIDE
1 | Pulse Oximeter | Nellcor | n-20 | 20964546 |  |  | 1 |  |  |  |  | Loose lead in electrode. Electrode replaced. Old electrode repaired | 1 |  | Endoscopy
2 | Pulse Oximeter | Biochem | 3040G | 5802170G |  |  |  | 1 |  |  |  | Only reads "E1" on screen. Believed to be an error with RAM according to manual. |  | 1 | MICU
3 | Pulse Oximeter | Biochem | 3040G | 5800617G |  |  | 1 |  |  |  |  | Functions sporadically but sometimes displays "E2." Solved by switching top circuit board with another machine. | 1 |  | MICU
4 | Patient Monitor | Agilent | M1770A | CND4751831 | [5 empty] | 1 |  | Too much noise in readings originally, but once more gel was used the readings were normal. | 1 |  | Endoscopy
5 | ECG Machine | Schiller | AT-2 plus | 025. 13111 |  |  |  | 1 |  |  |  | Keyboard pushed down into the mashine. Glued to fix | 1 |  | Endoscopy
6 | ECG Machine | Schiller | AT-2 plus | 025. 13111 | [5 empty] | 1 |  | Too much noise in readings originally, but once more gel was used the readings were normal. | 1 |  | Endoscopy
7 | X-Ray Machine | Philips | DuoDiagnost | 4558730/000458 | [5 empty] | 1 |  | Download of service manual and help troubleshooting |  | 1 | Radiology
8 | X-Ray Film View Box | Kodak | DirectView cr850 system | [6 empty] | 1 |  | Download of installation manual | 1 |  | Radiology
9 | Patient Monitor | Hewlett-packet | Omnicare Neonatal V24C | 3713A19808 |  |  | 1 |  |  |  |  | Reassembly of internal wires | 1 |  | MICU
10 | Other | Space heater | N/A | N/A |  |  | 1 |  |  |  |  | Rewirering and new plug | 1 |  | Engineering
11 | Aspirator/Suction Machine | Areosol Medical | B.S.5724 | GSF KCMC05675 |  |  |  | 1 |  |  |  | New filter, new tubing adn lubrication of motor | 1 |  | Surgical
12 | Oxygen Concentrator | Nuvo | Mark 5 | N/A | [6 empty] | 1 | Cleaned and tested for efficiency using flame test. | 1 |  | SICUA
13 | Patient Monitor | HP | 78352C | CN818 05721 | [6 empty] | 1 | Broken probe. Replaced | 1 |  | MICU
14 | Patient Monitor | HP | 78352C | 3111G04647 |  |  |  | 1 |  |  |  | Broken fuse and wrong transformer setting | 1 |  | MICU
15 | Printer | HP | 78173B | 2952800948 |  |  |  | 1 |  |  |  | Broken fuse and wrong transformer setting | 1 |  | MICU
16 | Patient Monitor | HP | V24CT | 4006a73100 |  |  |  | 1 |  |  |  | Replugging internal wire and note on longer startup time | 1 |  | SICU B
17 | Patient Monitor | HP | M1041A | 3604G16084 | 1 | [6 empty] | Replugging internal wire, cleaning power supply for dust | 1 |  | SICU B | 1997
18 | Patient Monitor | HP | V24CT | 3707A15810 |  |  |  | 1 |  |  |  | Replugging internal wire, cleaning power supply for dust | 1 |  | MICU
19 | Patient Monitor | HP | V24CT | 3642A15103 |  |  |  | 1 |  |  |  | Replugging internal wire, cleaning power supply for dust |  | 1 | MICU
20 | Patient Monitor | HP | V24CT | 3642A15081 |  |  |  | 1 |  |  |  | Replugging internal wire, cleaning power supply for dust |  | 1 | MICU
21 | Patient Monitor | HP | V24CT | 3950A68003 |  |  |  | 1 |  |  |  | Replugging internal wire, cleaning power supply for dust | 1 |  | SICU
22 | Patient Monitor | HP | V24CT | 3727A20688 |  |  |  | 1 |  |  |  | User error | 1 |  | SICU
23 | Patient Monitor | HP | V24CT | 3713A19808 |  |  |  | 1 |  |  |  | User error | 1 |  | SICU
24 | Patient Monitor | GE Dinamap | Pro1000 | WAA04340027SA |  |  | 1 |  |  |  |  | Missing power plug and dial. Doesn't turn on with power connected |  | 1 | MICU |  | Found at Engineering department
25 | Patient Monitor | GE Dinamap | Pro1000 | WAA04340033SA |  |  | 1 |  |  |  |  | Turns on. Missing probe |  | 1 | MICU
26 | X-Ray Film View Box | Kodak | Coldlight Illuminator | N/A | 1 | [6 empty] | Missing plug | 1 |  | OR 2
27 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
28 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
29 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
30 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
31 | Lamp, surgical | Skylux | ACE-1 8350 | A-2380 | 1 | [6 empty] | New fuse | 1 |  | OR 2
32 | Pulse Oximeter | Denk Pharma | Pocket edition | N/A |  |  | 1 |  |  |  |  | Probably personal device for a nurse. Corrosion between leads inside removed | 1 |  | Surgical
33 | Pulse Oximeter | MASimo | RAD/9 | A00127 | [6 empty] | 1 | Need of probe |  | 1 | SICU B | 2003 |  | GSF
34 | Pulse Oximeter | Datex | SATLITE TRANS | 402301 | [6 empty] | 1 | Working with probe | 1 |  | Surgical | 1993
35 | Pulse Oximeter | Datex | SATLITE TRANS | 403996 | [6 empty] | 1 | Functions but missing probe | 1 |  | Surgical | 1993
36 | Pulse Oximeter | Datex | SATLITE TRANS | 401057 | [6 empty] | 1 | Functions but missing probe | 1 |  | Surgical | 1993
37 | Pulse Oximeter | Datex | SATLITE TRANS | 400004 | [6 empty] | 1 | Functions but missing probe | 1 |  | Surgical | 1993
38 | Pulse Oximeter | INVIVO | 4500 PLUS 3 | 0X402978 | [6 empty] | 1 | Unsure of function because probe is missing |  | 1 | Surgical
39 | Pulse Oximeter | INVIVO | 4500 PLUS 3 | 0X400249 | [6 empty] | 1 | Unsure of function because probe is missing |  | 1 | Surgical
40 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | [5 empty] | 1 |  | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 |  | SICUB
41 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | [5 empty] | 1 |  | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 |  | SICUB
42 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | [5 empty] | 1 |  | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 |  | SICUB
43 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | [5 empty] | 1 |  | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 |  | SICUB
44 | Electrosurgery Unit (ESU) | Pfizer | Valleylab FORCE 40 | R4B4225S |  |  |  |  | 1 |  |  | Changed power cord. | 1 |  | OR3
45 | Electrosurgery Unit (ESU) | Pfizer | Valleylab FORCE 40 | R5I7616S |  |  |  |  | 1 |  |  | Power Button was difficult to press |  | 1 | OR1
45 | Anesthesia Machine | Dameca | 10590 | 9348051 | 1 | [6 empty] | Rewiring of tubing, patching bellow, reconfiguration of driving gass and training of staff | 1 |  | OR3
47 | Anesthesia Machine | Dameca | 10590 | 9348050 | 1 | [6 empty] | Rewiring of tubing, reconfiguration of driving gass and training of staff | 1 |  | OR4
48 | Ventilator | Royal medical | Vent-V | 4077 | 1 | [6 empty] | Tubing rewiring and training of staff | 1 |  | Anaesthesia room
49 | Capnograph | Datex | ULT-s-23-01 | 30134 |  |  | 1 |  |  |  |  | New power cord | 1 |  | OR 4
50 | Blood Pressure Device, Automatic (NIBP) | Unimed | N/A | N/A |  |  |  | 1 |  |  |  | Leaking cuff. Sealed with silicon. | 1 |  | SICUB
51 | Blood Pressure Device, Automatic (NIBP) | N/A | N/A | N/A |  |  |  | 1 |  |  |  | Replaced attachment nozzle. | 1 |  | Eye Dept
52 | Ventilator | Hand ventilator ballon |  | N/A | 1 | [6 empty] | Patch of holes | 1 |  | Eye Dept
53 | Ventilator | Hand ventilator ballon |  | N/A | 1 | [6 empty] | Patch of holes | 1 |  | ENT Dept
54 | Ventilator | Test lung |  | N/A | 1 | [6 empty] | Patch of holes | 1 |  | Eye Dept
55 | Anesthesia Machine | Siemens | Kion 50 | 190 | 1 | [6 empty] | Silicon of leakages | 1 |  | ENT Dept
56 | Anesthesia Machine | Carl Heyer | Narkomat M | 1375-91-348 |  |  |  | 1 |  |  |  | Support of Bellow and rewiring of internal tubing for need of only one cannister of oxygen | 1 |  | Eye Dept
57 | Anesthesia Machine | Acoma | PH-3f | 17774 | 1 | [6 empty] | rewiring of tubing and marking of connections | 1 |  | OR5 | 1995
58 | Ventilator | Acoma | ARF-900II | 6923 |  |  | 1 |  |  |  |  | resetting of controls |  | 1 | OR5 | 1995
59 | Anesthesia Machine | Acoma | PH-3f | 17779 | 1 | [6 empty] | rewiring of tubing and marking of connections | 1 |  | Hallway | 1995
60 | Autoclave (lab, surgery, and other) | London surgical equipment supplies ltd | sterilizer | N/A |  |  | 1 |  |  |  |  | resoldering of connections | 1 |  | Eye Dept
61 | Autoclave (lab, surgery, and other) | Matron autocalve | 520 | 1238 |  |  | 1 |  |  |  |  | Recalibration of pressure switch |  | 1 | Eye Dept
62 | Other | Nidek | KM-500 | 31944 |  |  |  |  | 1 |  |  | Auto Keratometer. Rewired rechargable battery pack and changed batteries | 1 |  | Eye Dept | 2009
63 | Defibrillator (automatic and manual) | BPL | DF 2509 | BZMA5L1333 |  |  |  |  | 1 |  |  | The defibrillator can only reach up to 20J, not the necessary 200J. Probably needs new capacitor. |  | 1 | MICU
64 | Defibrillator (automatic and manual) | Physio-Control | Lifepak 9P | 00007247 |  |  | 1 |  |  |  |  | The screen is jumpy and most buttons to not work. Unable to charge paddles |  | 1 | Recovery Room
65 | Defibrillator (automatic and manual) | Medical Data Electronics Inc. | MDE Escort | 1447 |  |  |  |  | 1 |  |  | Some button lights will turn on but the unit does not charge or display anything on the screen. |  | 1 | Surgical
66 | Defibrillator (automatic and manual) | Hewlett Packard | 43120A Defibrillator | 2612A51808 |  |  |  |  | 1 |  |  | Machine will not charge up. Batteries are dead and we were unable to find new rechargable ones of correct size and voltage. |  | 1 | Recovery Room
67 | Other | Humidifier |  | N/A |  |  |  | 1 |  |  |  | Filled leaks with silicon and reassembled parts. | 1 |  | Surgical
68 | Other | Humidifier |  | N/A |  |  |  | 1 |  |  |  | Filled leaks with silicon and reassembled parts. | 1 |  | Surgical
69 | Other | Humidifier |  | N/A |  |  |  | 1 |  |  |  | Filled leaks with silicon and reassembled parts. | 1 |  | Surgical
70 | Defibrillator (automatic and manual) | Nihon Kohden Coorporation | Cardiolife TEC-710K | 20416 |  |  |  |  | 1 |  |  | New power cord | 1 |  | MICU
71 | Anesthesia Machine | Siemens | Kion 50 | 200 | 1 | [6 empty] | Blokage of tubing | 1 |  | ENT Dept
72 | Ventilator | Royal medical | Vent-V | 4074 | 1 | [6 empty] | Tubing rewiring and training of staff | 1 |  | Anaesthesia room
73 | Ventilator | Royal medical | Vent-V | 4079 | 1 | [6 empty] | Tubing rewiring and training of staff |  | 1 | Anaesthesia room
74 | Vacuum Extractor (for delivery) | Drager | Evac40 | 170032181-073061 | [5 empty] | 1 |  | Suction pump with user error. The staff did not know that the machine can be turned off using the foot pedal. | 1 |  |  | 2010
75 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
76 | Other | Humidifier | N/A | N/A |  |  |  | 1 |  |  |  | Leaking. Silicon and new parts | 1 |  | Surgical
77 | Anesthesia Machine | Ohmeda | Modulus II Plus | ABQS01087 |  |  |  |  | 1 |  |  | Replacement of broken component in power supply and Stepdown transformer | 1 |  | OR3
78 | Nebulizer | Fisher and paykel | MR4010 | 9641jhb13786 | 1 | [6 empty] | Tested settings added typing | 1 |  | Anaesthesia room
79 | Autoclave (lab, surgery, and other) | Melag | Euroclave 23v-s | 0523v-s1484 | 1 | [6 empty] | Error code for leakage. Leakage not fund |  | 1 | Eye Dept
80 | Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable) | Welch Allyn | N/A | 50726213424 |  |  |  | 1 |  |  |  | Leakage in the tubing. Replaced. | 1 |  | Urology Department
81 | Other | Vaporizer by General Anaesthetic Services | Unit F2B | 00 6609 |  |  |  | 1 |  |  |  | Used ethanol and oil to loosen dial. Took apart and cleaned. Risked lives by puposely exposing ourselves to halothane. | 1 |  | Surgical
82 | Other | Vaporizer by Cyprane Keighley | FLUOTEC3 | 389689 |  |  |  | 1 |  |  |  | Used ethanol and oil to loosen dial. Took apart and cleaned. Risked lives by puposely exposing ourselves to halothane. | 1
83 | Other | Voltage Regulator by AVR | AS-500-23E | 13 |  |  |  |  | 1 |  |  | Changed power cord. | 1 |  | Surgical | 1995
84
85
86
87
88 | [16 empty] | 83
89
90
91
92
93
94
95
96
97
98
99
100

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Nebulizer
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Patient Monitor
Photocopier
Phototherapy device
Printer
Projector
Pulse Oximeter
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Ventilator
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Other

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
KCMC Moshi14. July 20151 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 1 | Pulse Oximeter | Nellcor | n-20 | 20964546 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Loose lead in electrode. Electrode replaced. Old electrode repaired | 1 | 0
KCMC Moshi14. July 20152 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 2 | Pulse Oximeter | Biochem | 3040G | 5802170G | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Only reads "E1" on screen. Believed to be an error with RAM according to manual. | 0 | 1
KCMC Moshi14. July 20153 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 3 | Pulse Oximeter | Biochem | 3040G | 5800617G | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Functions sporadically but sometimes displays "E2." Solved by switching top circuit board with another machine. | 1 | 0
KCMC Moshi14. July 20154 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 4 | Patient Monitor | Agilent | M1770A | CND4751831 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Too much noise in readings originally, but once more gel was used the readings were normal. | 1 | 0
KCMC Moshi14. July 20155 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 5 | ECG Machine | Schiller | AT-2 plus | 025. 13111 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Keyboard pushed down into the mashine. Glued to fix | 1 | 0
KCMC Moshi14. July 20156 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 6 | ECG Machine | Schiller | AT-2 plus | 025. 13111 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Too much noise in readings originally, but once more gel was used the readings were normal. | 1 | 0
KCMC Moshi14. July 20157 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 7 | X-Ray Machine | Philips | DuoDiagnost | 4558730/000458 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Download of service manual and help troubleshooting | 0 | 1
KCMC Moshi14. July 20158 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 8 | X-Ray Film View Box | Kodak | DirectView cr850 system | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Download of installation manual | 1 | 0
KCMC Moshi14. July 20159 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 9 | Patient Monitor | Hewlett-packet | Omnicare Neonatal V24C | 3713A19808 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Reassembly of internal wires | 1 | 0
KCMC Moshi14. July 201510 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 10 | Other | Space heater | N/A | N/A | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Rewirering and new plug | 1 | 0
KCMC Moshi14. July 201511 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 11 | Aspirator/Suction Machine | Areosol Medical | B.S.5724 | GSF KCMC05675 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | New filter, new tubing adn lubrication of motor | 1 | 0
KCMC Moshi14. July 201512 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 12 | Oxygen Concentrator | Nuvo | Mark 5 | N/A | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaned and tested for efficiency using flame test. | 1 | 0
KCMC Moshi14. July 201513 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 13 | Patient Monitor | HP | 78352C | CN818 05721 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Broken probe. Replaced | 1 | 0
KCMC Moshi14. July 201514 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 14 | Patient Monitor | HP | 78352C | 3111G04647 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Broken fuse and wrong transformer setting | 1 | 0
KCMC Moshi14. July 201515 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 15 | Printer | HP | 78173B | 2952800948 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Broken fuse and wrong transformer setting | 1 | 0
KCMC Moshi14. July 201516 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 16 | Patient Monitor | HP | V24CT | 4006a73100 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replugging internal wire and note on longer startup time | 1 | 0
KCMC Moshi14. July 201517 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 17 | Patient Monitor | HP | V24CT | 3707A15810 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replugging internal wire, cleaning power supply for dust | 1 | 0
KCMC Moshi14. July 201518 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 18 | Patient Monitor | HP | V24CT | 3642A15103 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replugging internal wire, cleaning power supply for dust | 0 | 1
KCMC Moshi14. July 201519 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 19 | Patient Monitor | HP | V24CT | 3642A15081 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replugging internal wire, cleaning power supply for dust | 0 | 1
KCMC Moshi14. July 201520 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 20 | Patient Monitor | HP | V24CT | 3950A68003 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replugging internal wire, cleaning power supply for dust | 1 | 0
KCMC Moshi14. July 201521 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 21 | Patient Monitor | HP | V24CT | 3727A20688 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | User error | 1 | 0
KCMC Moshi14. July 201522 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 22 | Patient Monitor | HP | V24CT | 3713A19808 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | User error | 1 | 0
KCMC Moshi14. July 201523 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 23 | Patient Monitor | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
KCMC Moshi14. July 201524 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 24 | Patient Monitor | GE Dinamap | Pro1000 | WAA04340027SA | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Missing power plug and dial. Doesn't turn on with power connected | 0 | 1
KCMC Moshi14. July 201525 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 25 | Patient Monitor | GE Dinamap | Pro1000 | WAA04340033SA | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Turns on. Missing probe | 0 | 1
KCMC Moshi14. July 201526 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 26 | X-Ray Film View Box | Kodak | Coldlight Illuminator | N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Missing plug | 1 | 0
KCMC Moshi14. July 201527 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 27 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201528 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 28 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201529 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 29 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201530 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 30 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201531 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 31 | Lamp, surgical | Skylux | ACE-1 8350 | A-2380 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | New fuse | 1 | 0
KCMC Moshi14. July 201532 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 32 | Pulse Oximeter | Denk Pharma | Pocket edition | N/A | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Probably personal device for a nurse. Corrosion between leads inside removed | 1 | 0
KCMC Moshi14. July 201533 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 33 | Pulse Oximeter | MASimo | RAD/9 | A00127 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Need of probe | 0 | 1
KCMC Moshi14. July 201534 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 34 | Pulse Oximeter | Datex | SATLITE TRANS | 402301 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Working with probe | 1 | 0
KCMC Moshi14. July 201535 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 35 | Pulse Oximeter | Datex | SATLITE TRANS | 403996 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Functions but missing probe | 1 | 0
KCMC Moshi14. July 201536 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 36 | Pulse Oximeter | Datex | SATLITE TRANS | 401057 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Functions but missing probe | 1 | 0
KCMC Moshi14. July 201537 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 37 | Pulse Oximeter | Datex | SATLITE TRANS | 400004 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Functions but missing probe | 1 | 0
KCMC Moshi14. July 201538 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 38 | Pulse Oximeter | INVIVO | 4500 PLUS 3 | 0X402978 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Unsure of function because probe is missing | 0 | 1
KCMC Moshi14. July 201539 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 39 | Pulse Oximeter | INVIVO | 4500 PLUS 3 | 0X400249 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Unsure of function because probe is missing | 0 | 1
KCMC Moshi14. July 201540 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 40 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | 0 | 0 | 0 | 0 | 0 | 1 | 0 | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 | 0
KCMC Moshi14. July 201541 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 41 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | 0 | 0 | 0 | 0 | 0 | 1 | 0 | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 | 0
KCMC Moshi14. July 201542 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 42 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | 0 | 0 | 0 | 0 | 0 | 1 | 0 | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 | 0
KCMC Moshi14. July 201543 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 43 | Blood Pressure Device, Automatic (NIBP) | Unimed | U1880S | N/A | 0 | 0 | 0 | 0 | 0 | 1 | 0 | User Error. Reconnected to machine with correct adapter and tested cuffs | 1 | 0
KCMC Moshi14. July 201544 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 44 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
KCMC Moshi14. July 201545 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 45 | Electrosurgery Unit (ESU) | Pfizer | Valleylab FORCE 40 | R4B4225S | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Changed power cord. | 1 | 0
KCMC Moshi14. July 201545 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 45 | Electrosurgery Unit (ESU) | Pfizer | Valleylab FORCE 40 | R5I7616S | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Power Button was difficult to press | 0 | 1
KCMC Moshi14. July 201547 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 47 | Anesthesia Machine | Dameca | 10590 | 9348051 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Rewiring of tubing, patching bellow, reconfiguration of driving gass and training of staff | 1 | 0
KCMC Moshi14. July 201548 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 48 | Anesthesia Machine | Dameca | 10590 | 9348050 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Rewiring of tubing, reconfiguration of driving gass and training of staff | 1 | 0
KCMC Moshi14. July 201549 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 49 | Ventilator | Royal medical | Vent-V | 4077 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Tubing rewiring and training of staff | 1 | 0
KCMC Moshi14. July 201550 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 50 | Capnograph | Datex | ULT-s-23-01 | 30134 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | New power cord | 1 | 0
KCMC Moshi14. July 201551 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 51 | Blood Pressure Device, Automatic (NIBP) | Unimed | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking cuff. Sealed with silicon. | 1 | 0
KCMC Moshi14. July 201552 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 52 | Blood Pressure Device, Automatic (NIBP) | N/A | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replaced attachment nozzle. | 1 | 0
KCMC Moshi14. July 201553 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 53 | Ventilator | Hand ventilator ballon | 0 | N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Patch of holes | 1 | 0
KCMC Moshi14. July 201554 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 54 | Ventilator | Hand ventilator ballon | 0 | N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Patch of holes | 1 | 0
KCMC Moshi14. July 201555 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 55 | Ventilator | Test lung | 0 | N/A | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Patch of holes | 1 | 0
KCMC Moshi14. July 201556 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 56 | Anesthesia Machine | Siemens | Kion 50 | 190 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Silicon of leakages | 1 | 0
KCMC Moshi14. July 201557 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 57 | Anesthesia Machine | Carl Heyer | Narkomat M | 1375-91-348 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Support of Bellow and rewiring of internal tubing for need of only one cannister of oxygen | 1 | 0
KCMC Moshi14. July 201558 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 58 | Anesthesia Machine | Acoma | PH-3f | 17774 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | rewiring of tubing and marking of connections | 1 | 0
KCMC Moshi14. July 201559 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 59 | Ventilator | Acoma | ARF-900II | 6923 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | resetting of controls | 0 | 1
KCMC Moshi14. July 201560 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 60 | Anesthesia Machine | Acoma | PH-3f | 17779 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | rewiring of tubing and marking of connections | 1 | 0
KCMC Moshi14. July 201561 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 61 | Autoclave (lab, surgery, and other) | London surgical equipment supplies ltd | sterilizer | N/A | 0 | 0 | 1 | 0 | 0 | 0 | 0 | resoldering of connections | 1 | 0
KCMC Moshi14. July 201562 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 62 | Autoclave (lab, surgery, and other) | Matron autocalve | 520 | 1238 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Recalibration of pressure switch | 0 | 1
KCMC Moshi14. July 201563 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 63 | Other | Nidek | KM-500 | 31944 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Auto Keratometer. Rewired rechargable battery pack and changed batteries | 1 | 0
KCMC Moshi14. July 201564 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 64 | Defibrillator (automatic and manual) | BPL | DF 2509 | BZMA5L1333 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | The defibrillator can only reach up to 20J, not the necessary 200J. Probably needs new capacitor. | 0 | 1
KCMC Moshi14. July 201565 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 65 | Defibrillator (automatic and manual) | Physio-Control | Lifepak 9P | 00007247 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | The screen is jumpy and most buttons to not work. Unable to charge paddles | 0 | 1
KCMC Moshi14. July 201566 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 66 | Defibrillator (automatic and manual) | Medical Data Electronics Inc. | MDE Escort | 1447 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Some button lights will turn on but the unit does not charge or display anything on the screen. | 0 | 1
KCMC Moshi14. July 201567 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 67 | Defibrillator (automatic and manual) | Hewlett Packard | 43120A Defibrillator | 2612A51808 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Machine will not charge up. Batteries are dead and we were unable to find new rechargable ones of correct size and voltage. | 0 | 1
KCMC Moshi14. July 201568 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 68 | Other | Humidifier | 0 | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Filled leaks with silicon and reassembled parts. | 1 | 0
KCMC Moshi14. July 201569 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 69 | Other | Humidifier | 0 | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Filled leaks with silicon and reassembled parts. | 1 | 0
KCMC Moshi14. July 201570 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 70 | Other | Humidifier | 0 | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Filled leaks with silicon and reassembled parts. | 1 | 0
KCMC Moshi14. July 201571 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 71 | Defibrillator (automatic and manual) | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0 | 1 | 0 | 0 | New power cord | 1 | 0
KCMC Moshi14. July 201572 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 72 | Anesthesia Machine | Siemens | Kion 50 | 200 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Blokage of tubing | 1 | 0
KCMC Moshi14. July 201573 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 73 | Ventilator | Royal medical | Vent-V | 4074 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Tubing rewiring and training of staff | 1 | 0
KCMC Moshi14. July 201574 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 74 | Ventilator | Royal medical | Vent-V | 4079 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Tubing rewiring and training of staff | 0 | 1
KCMC Moshi14. July 201575 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 75 | Vacuum Extractor (for delivery) | Drager | Evac40 | 170032181-073061 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Suction pump with user error. The staff did not know that the machine can be turned off using the foot pedal. | 1 | 0
KCMC Moshi14. July 201576 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 76 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201577 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 77 | Other | Humidifier | N/A | N/A | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leaking. Silicon and new parts | 1 | 0
KCMC Moshi14. July 201578 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 78 | Anesthesia Machine | Ohmeda | Modulus II Plus | ABQS01087 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Replacement of broken component in power supply and Stepdown transformer | 1 | 0
KCMC Moshi14. July 201579 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 79 | Nebulizer | Fisher and paykel | MR4010 | 9641jhb13786 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Tested settings added typing | 1 | 0
KCMC Moshi14. July 201580 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 80 | Autoclave (lab, surgery, and other) | Melag | Euroclave 23v-s | 0523v-s1484 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Error code for leakage. Leakage not fund | 0 | 1
KCMC Moshi14. July 201581 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 81 | Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable) | Nihon Kohden Coorporation | Cardiolife TEC-710K | 20416 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leakage in the tubing. Replaced. | 1 | 0
KCMC Moshi14. July 201582 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 82 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
KCMC Moshi14. July 201583 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 83 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
KCMC Moshi14. July 201584 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 84 | Other | Vaporizer by General Anaesthetic Services | 00 6609 | #REF! | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Used ethanol and oil to loosen dial. Took apart and cleaned. Risked lives by puposely exposing ourselves to halothane. | 1 | 0
KCMC Moshi14. July 201585 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 85 | Other | Vaporizer by Cyprane Keighley | FLUOTEC3 | 389689 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Used ethanol and oil to loosen dial. Took apart and cleaned. Risked lives by puposely exposing ourselves to halothane. | 1 | 0
KCMC Moshi14. July 201586 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 86 | Other | Voltage Regulator by AVR | AS-500-23E | 13 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Changed power cord. | 1 | 0
KCMC Moshi14. July 201587 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 87 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201588 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 88 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201589 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201590 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201591 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201592 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201593 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 93 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201594 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 94 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201595 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201596 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 96 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201597 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 97 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201598 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 98 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 201599 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
KCMC Moshi14. July 2015100 | KCMC Moshi | 14. July 2015 | Amy Xiong, Lauren Barnes, Morten Lindhardt Madsen | Tanzania | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0