## Work Summary Form
[5 empty] | Engineering World Health | [8 empty] | The Prizery, Suite 200
302 East Pettigrew Street
Durham, NC 27701
www.ewh.org
Contacto: Inka Johnson, inka@ewh.org
|  |  | DATE | 20/07/2015 |  |  | ENGINEER NAME(S) |  |  |  |  | Deanna Hood, Sindora Baddam
|  |  | COUNTRY | Nicaragua |  |  | Hospital/CITY |  |  |  |  | Asunción, Juigalpa and surrounding health centres in Chontales
| EQUIPMENT INFORMATION | [5 empty] | PROBLEM AND SOLUTION | [7 empty] | RESULTS
| Location | Date majority of work was done | Equipment type | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
1 | Hospital Asunción, Juigalpa | 2015-06-17 | Incubator (infant) | Dräger | 8000 SC / 2m 20315 | ARDF-00 0 1 |  |  |  | 1 |  |  |  | Perspex door fractured. Maintenance staff made a metal brace for door and we did preventative maintenance on the internals. Fixed a buzzer that had come apart and cable holder that was otherwise not used. | 1
2 | Hospital Asunción, Juigalpa | 2015-06-18 | Infant Warmer (radiant or other) | [7 empty] | 1 |  |  | Active wire disconnected at the plug; spliced with another connector. | 1
3 | Hospital Asunción, Juigalpa | 2015-06-18 | Ultrasound physiotherapy | Exogen 2000+ | 1408275 | 904076426 |  |  |  |  | 1 |  |  | Error 32 even with coupling gel. Must have meant low internal batteries (5x 3.6V lithium) because when using external power supply went away. Replaced batteries with 2x 9V alkaline. | 1
4 | Hospital Asunción, Juigalpa | 2015-06-19 | Capillary centrifuge | Chem index inc | DSC-030MH | 9603011-80 |  |  |  | 1 |  |  |  | Rubber barrier preventing vials from hitting outer wall missing. Replaced with a piece truck tyre tube. Cleaned out wasps nests from inside. Fixed support leg of motor so spinning disk doesn't scrape. | 1
5 | Hospital Asunción, Juigalpa | 2015-06-19 | Transformer | Chicago Digital Power | BAVR1006 |  |  |  | 1 |  |  |  |  | Power switch broken on surge protection box. Replaced with a switch from another. | 1
6 | Hospital Asunción, Juigalpa | 2015-06-21 | Blood Pressure Device, Automatic (NIBP) | Homedics | BPS-420WGN | 8020766 |  |  |  | 1 |  |  |  | Leak in cuff. Sealed and restitched edges of covering fabric. | 1
7 | Hospital Asunción, Juigalpa | 2015-06-21 | Ventilador/Respirador | AVEA | 17312-06 | BAV06802 | [7 empty] | Not ventilating at requested (or any consistent) rate. Gives errors of too much and too little pressure simultaneously. Gives errors of the circuit being occluded and disconnected simulatneously. Has been recalled by manufacturer for faulty pressure transducer. |  | 1
8 | Centro de Salud, Santo Tomás | 2015-06-23 | Capillary centrifuge | LW Scientific | LWS-M24 | 409234 |  |  |  | 1 |  |  |  | Replaced rubber barrier preventing vials from breaking. | 1
9 | Centro de Salud, Santo Tomás | 2015-06-23 | Capillary centrifuge | Chem index inc | DSC-030MH |  |  | 1 |  | 1 | 1 |  |  | Resoldered a wire burnt out near the motor. Replaced the cable with one with three pins because earth was missing. Replaced rubber barrier preventing vials from breaking. | 1
10 | Centro de Salud, Santo Tomás | 2015-06-23 | Centrifuge |  | LWS-M24 combo | 0211180 |  |  |  | 1 |  | 1 |  | Unbalanced load biggest problem, also needed an adjustment of door closed sensor. | 1
11 | Centro de Salud, Santo Tomás | 2015-06-23 | Scales (laboratory and in wards) | ADE | M10813 | 101002625 |  |  |  | 1 |  |  |  | Not balancing on necessary points correctly, and clip missing in weighing mechanism replaced with cable tie. | 1
12 | Centro de Salud, Santo Tomás | 2015-06-24 | Microscope | LW Scientific | Revelation III | 003344 | [5 empty] | 1 |  | New bulb blew once put in microscope - using 6V bulb in 12V microscope. Clearly labelled now. | 1
13 | Centro de Salud, Santo Tomás | 2015-06-23 | Nebulizer | Medel | AG Classic | 034819 |  |  |  | 1 |  |  |  | Piston disconnected from motor. Secured with cable tie to prevent coming off in the future. | 1
14 | Centro de Salud, Santo Tomás | 2015-06-23 | Nebulizer | Omron | NE-C801 | 20110100194UF | [6 empty] | 1 | Not making enough output - no internal faults found, but very lightweight compared to other nebulizers. |  | 1
15 | Centro de Salud, Santo Tomás | 2015-06-23 | Nebulizer | Shuco | S3000 | 020700045670 | [5 empty] | 1 |  | Old and loud but functioning. | 1
16 | Centro de Salud, Santo Tomás | 2015-06-23 | Nebulizer | Shuco | S3000 |  |  |  |  | 1 |  |  |  | Tubing glued to output pipe for some reason has a kink. When removed, output pipe broke off. |  | 1
17 | Centro de Salud, Santo Tomás | 2015-06-23 | Fetal steth (fetoscope or Doppler) | Huntleigh | SonicaidOne | 005794 | [6 empty] | 1 | Readings are wrong. No faults identified. |  | 1
18 | Centro de Salud, Santo Tomás | 2015-06-24 | Capillary centrifuge | Miami medical |  | 27743 |  |  |  | 1 |  |  |  | Rubber barrier preventing vials from hitting outer wall missing. Replaced with a piece truck tyre tube. | 1 |  | *reworked on it later due to rubber barrier being too high preventing the top from closing
19 | Centro de Salud, Santo Tomás | 2015-06-26 | Microscope | Olympus | MX31RBSFA | 3C044074 |  |  | 1 |  |  |  |  | Needed bulb socket connection soldered back on. | 1
20 | Centro de Salud, Santo Tomás | 2015-06-25 | Oven (laboratory, not kitchen) | Memmert |  | 793511 |  |  | 1 |  |  |  |  | Burnt cables re-soldered. | 1
21 | Centro de Salud, Santo Tomás | 2015-06-25 | Cell counter | Intramedica | MG-4318 |  |  |  |  | 1 |  |  |  | Total not incrementing for all keys. Cable tie on an axle stopped the axle lifting too far from a worn hole. Broken springs replaced with rubber bands. | 1
22 | Centro de Salud, Santo Tomás | 2015-06-25 | Cell counter | Intramedica | MG-4318 |  |  |  |  | 1 |  |  |  | Broken springs replaced with rubber bands. | 1
23 | Centro de Salud, Santo Tomás | 2015-06-26 | Microscope | Olympus | MX31RBSFA | 3D09715 |  |  |  |  | 1 |  |  | No voltage at bulb socket. Ran out of time to investigate further. |  | 1
24 | Centro de Salud, Santo Tomás | 2015-06-26 | Microscope | Olympus | MX31RBSFA | [5 empty] | 1 |  |  | No voltage at bulb socket. Ran out of time to investigate further. |  | 1
25 | Centro de Salud, Santo Tomás | 2015-06-26 | Microscope | Askmania/SLG Chemnitz | RME5 | 1193174 |  |  |  | 1 |  |  |  | Labelled specs for missing bulb, fixed hinge. | 1
26 | Centro de Salud, Santo Tomás | 2015-06-26 | Scales (laboratory and in wards) | [6 empty] | 1 |  |  |  | Needed a cable tie to secure bassinet to scale. | 1
27 | Centro de Salud, La Libertad | 2015-06-29 | Fetal steth (fetoscope or Doppler) | Biocare | FM-200 | 3314040595 |  |  | 1 |  |  |  |  | Broken speaker but working with audio jack output. | 1
28 | Centro de Salud, La Libertad | 2015-06-29 | Exercise bicycle | Exercycle | H8705M | S6117539 |  |  |  |  | 1 |  |  | Repaired loose cables in original power jack. | 1
29 | Centro de Salud, La Libertad | 2015-06-29 | Physiotherapy electrodes | [5 empty] | 1 |  |  |  |  | Repaired loose connections between electrode pads and cables with epoxy. | 1
30 | Centro de Salud, La Libertad | 2015-06-29 | Spectrophotometer/Colorimeter | 2MM | Diagnostiko-550 | 4505-1001 |  |  | 1 |  |  |  |  | Needed a new bulb (quite complicated to replace). Printer had been turned off. | 1
31 | Centro de Salud, La Libertad | 2015-06-29 | Microscope | LW Scientific | Revelation III | 128465 |  |  |  | 1 |  |  |  | Slide-holding mechanism was off its rails and bent out of shape. | 1
32 | Hospital Asunción, Juigalpa | 2015-06-30 | Patient monitor | Inohero | H8 | [8 empty] | Poor connection on pulse oximeter socket. Didn't find time to return to work on it. |  | 1
33 | Hospital Asunción, Juigalpa | 2015-06-30 | Patient monitor | Goldway | UT4000F | CN4FAAJW00562 | [7 empty] | Bad pulse oximeter output from socket. Poor ecg signal. Didn't find time to return to work on it. |  | 1
34 | Hospital Asunción, Juigalpa | 2015-06-30 | Patient monitor | Goldway | UT4000F | CN4FABAQ00624 |  |  | 1 |  |  |  |  | Bad pulse oximeter output from socket. Circuits and connections are all good, perhaps chips damaged by static electricity. Referred hospital to local servicer before leaving. |  | 1
35 | Hospital Asunción, Juigalpa | 2015-06-30 | Ventilator | Neumovent | graphnet advance | 120710341081A1V | 1 | [6 empty] | Leak in filter for pressurised air input. Didn't find time to return to work on it. |  | 1
36 | Centro de Salud, Acoyapa | 2015-07-06 | Centrifuge | Clay Adams | Compact II | 14-04-4-59-3 |  |  |  | 1 |  |  |  | Lid broken off, rejoined with epoxy. Seal missing, replaced with IV tubing. | 1
37 | Centro de Salud, Acoyapa | 2015-07-06 | Shaker Machine (laboratory) | Chem index inc | PS600T | 831260 |  |  |  | 1 |  |  |  | Pipette holder disconnected from shaking rod (missing screw), rejoined with with a ring of rubber tubing. | 1
38 | Centro de Salud, Acoyapa | 2015-07-06 | Oven (laboratory, not kitchen) | Rajas Scientific | Dry Heat Sterilizer | None | [5 empty] | 1 |  | 220V UK plug. Missing manual and unknown how to change desired temperature. |  | 1
39 | Centro de Salud, Acoyapa | 2015-07-06 | Nebulizer | Wellpro | NBA-02WD | 356 |  |  |  |  | 1 |  |  | Fried inside. |  | 1
40 | Centro de Salud, Acoyapa | 2015-07-06 | Nebulizer | Wellpro | NBA-02WD | 355 | [5 empty] | 1 |  | Working fine. Advised that manual says to give it half hour breaks after half an hour of use. | 1
41 | Centro de Salud, Acoyapa | 2015-07-07 | Electrosurgery Unit (ESU) | Birtcher | Hyfrecator 732 | B27443 |  |  | 1 |  |  |  |  | Only partially works/some of the time. Loose connection inside the pen. | 1
42 | Centro de Salud, Acoyapa | 2015-07-07 | Oven (laboratory, not kitchen) | Memmert | U10? | 14-04-4-101-1 | [5 empty] | 1 |  | Were told it wasn't regulating the temperature properly and that it burnt paper covered equipment. Seems fine but temperature dial needed to be recalibrated and digital thermometer readings are incorrect. | 1
43 | Centro de Salud, Acoyapa | 2015-07-07 | Nebulizer | DeVILBISS | PulmoAide 5650D | D4092803 | [5 empty] | 1 |  | Working fine. Made a hole so case lid can be closed will cord is attached to not block air filter. | 1
44 | Centro de Salud, Acoyapa | 2015-07-07 | Nebulizer | Drive Medical | Spider Neb II | M70308730 |  |  |  | 1 |  |  |  | Piston disconnected from motor. Secured with cable tie to prevent coming off in the future. | 1
45 | Centro de Salud, Acoyapa | 2015-07-07 | Nebulizer | John Bunn | Neb-u-Lite | AJM296RP6160 |  |  |  | 1 |  |  |  | Piston broken, replaced with the one from Item 39. | 1
46 | Laboratorio Regional, Juigalpa | 2015-07-08 | Oven (laboratory, not kitchen) | FischerScientific | Isotemp 506G | 60800197 | [5 empty] | 1 |  | Bothered that the numbers on the dial didn't represent the temperature settings. No changes made. | 1
47 | Laboratorio Regional, Juigalpa | 2015-07-09 | Rotator (laboratory) | Chem index inc | DSR 2100V | 831130153 |  |  |  | 1 |  |  |  | Broken belt, replaced with a ring of a blood pressure cuff. | 1
48 | Laboratorio Regional, Juigalpa | 2015-07-08 | Spectrophotometer/Colorimeter | Awareness Technology inc | Stat Fax 4542 | 4503-1134 | [5 empty] | 1 |  | Wasn't printing the calculations they wanted, changed the settings with them. | 1
49 | Laboratorio Regional, Juigalpa | 2015-07-08 | Capillary centrifuge | Intramedica | Illegible | Illegible |  |  |  | 1 |  |  |  | Rubber barrier preventing vials from hitting outer wall was old, replaced with a piece truck tyre tube. Spliced cable with one having three prongs. Replaced bypassed fuse. | 1
50 | Laboratorio Regional, Juigalpa | 2015-07-08 | Pipette | DragonLab | Multi 50-300uL | DX75176 |  |  |  | 1 |  |  |  | Button reconnected with epoxy. | 1
51 | Laboratorio Regional, Juigalpa | 2015-07-10 | Centrifuge | Chem index inc | DSC1512T | 2010310-22 | [6 empty] | 1 | Working but stirring up rust particles. As much corrosion as possible covered with epoxy. Fuse changed to correct value (3A to 5A). | 1
52 | Laboratorio Regional, Juigalpa | 2015-07-10 | Oven (laboratory, not kitchen) | Memmert | 530 | 660745 |  |  |  | 1 |  |  |  | Handle missing. Replaced with wire which hooked onto more wire on the side of the oven. | 1
53 | Laboratorio Regional, Juigalpa | 2015-07-10 | Oven (laboratory, not kitchen) | Lab-Line | 125 | 0198-0344 |  |  |  | 1 |  |  |  | Door clip damaged. Epoxied a hook and attached wire to the handle to act as a lock. | 1
54 | Laboratorio Regional, Juigalpa | 2015-07-10 | Spectrophotometer/Colorimeter | Rayto | RT-2100C | 451218015 | [5 empty] | 1 |  | Wanted a printer (HP D2660) connected to it but it only accepts five particular models. |  | 1
55 | Hospital Asunción, Juigalpa | 2015-07-15 | ECG Machine | Cardioline | ar600adv | 05081123 | [5 empty] | 1 |  | Somehow the language changed into english. The manual doesn't say how to change language and the manufacturer did not reply. Printed a translation sheet. | 1

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Cell counter
Centrifuge
Capillary centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Exercise bicycle
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Nebulizer
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Patient monitor
Photocopier
Phototherapy device
Physiotherapy electrodes
Pipette
Printer
Projector
Pulse Oximeter
Rotator (laboratory)
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Ventilator
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Otro
Ventilador/Respirador
Monitor de Paciente
Capnógrafo
Máquina de Anestesia
Oxímetro de pulso
Bomba de Infusión
Desfibrilador
Báscula de peso
sistema de cámara de vídeo
torniquete inflable
Calentador infantil
Cama ajustable
lavadora de botellas para los bebés
Calentador de líquidos
cortadora de pan
Horno para pan
máquina de tortillas
máquina de diálisis (en el sótano)
Ultrasound physiotherapy

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment type | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20151 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 1 | Incubator (infant) | Dräger | 8000 SC / 2m 20315 | ARDF-00 0 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Perspex door fractured. Maintenance staff made a metal brace for door and we did preventative maintenance on the internals. Fixed a buzzer that had come apart and cable holder that was otherwise not used. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20152 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 2 | Infant Warmer (radiant or other) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Active wire disconnected at the plug; spliced with another connector. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20153 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 3 | Ultrasound physiotherapy | Exogen 2000+ | 1408275 | 904076426 | 0 | 0 |  | 0 | 1 | 0 | 0 | Error 32 even with coupling gel. Must have meant low internal batteries (5x 3.6V lithium) because when using external power supply went away. Replaced batteries with 2x 9V alkaline. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20154 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 4 | Capillary centrifuge | Chem index inc | DSC-030MH | 9603011-80 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Rubber barrier preventing vials from hitting outer wall missing. Replaced with a piece truck tyre tube. Cleaned out wasps nests from inside. Fixed support leg of motor so spinning disk doesn't scrape. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20155 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 5 | Transformer | Chicago Digital Power | BAVR1006 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Power switch broken on surge protection box. Replaced with a switch from another. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20156 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 6 | Blood Pressure Device, Automatic (NIBP) | Homedics | BPS-420WGN | 8020766 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Leak in cuff. Sealed and restitched edges of covering fabric. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20158 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 8 | Capillary centrifuge | LW Scientific | LWS-M24 | 409234 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Replaced rubber barrier preventing vials from breaking. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20159 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 9 | Scales (laboratory and in wards) | ADE | M10813 | 101002625 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | Resoldered a wire burnt out near the motor. Replaced the cable with one with three pins because earth was missing. Replaced rubber barrier preventing vials from breaking. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201510 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 10 | Centrifuge | 0 | LWS-M24 combo | 0211180 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | Unbalanced load biggest problem, also needed an adjustment of door closed sensor. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201511 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 11 | #REF! | #REF! | #REF! | #REF! | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Not balancing on necessary points correctly, and clip missing in weighing mechanism replaced with cable tie. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201512 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 12 | Microscope | LW Scientific | Revelation III | 003344 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | New bulb blew once put in microscope - using 6V bulb in 12V microscope. Clearly labelled now. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201513 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 13 | Nebulizer | Shuco | S3000 | 020700045670 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Piston disconnected from motor. Secured with cable tie to prevent coming off in the future. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201514 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 14 | Nebulizer | Omron | NE-C801 | 20110100194UF | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Not making enough output - no internal faults found, but very lightweight compared to other nebulizers. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201515 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 15 | Nebulizer | #REF! | #REF! | #REF! | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Old and loud but functioning. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201516 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 16 | Nebulizer | Shuco | S3000 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Tubing glued to output pipe for some reason has a kink. When removed, output pipe broke off. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201517 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 17 | Fetal steth (fetoscope or Doppler) | Huntleigh | SonicaidOne | 005794 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Readings are wrong. No faults identified. | 0 | 1
#REF! | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201518 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 18 | Capillary centrifuge | Miami medical | 0 | 27743 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Rubber barrier preventing vials from hitting outer wall missing. Replaced with a piece truck tyre tube. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201519 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 19 | Microscope | Olympus | MX31RBSFA | 3C044074 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Needed bulb socket connection soldered back on. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201520 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 20 | Oven (laboratory, not kitchen) | Memmert | 0 | 793511 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Burnt cables re-soldered. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201523 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 23 | Microscope | Olympus | MX31RBSFA | 3D09715 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | No voltage at bulb socket. Ran out of time to investigate further. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201524 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 24 | Microscope | Olympus | MX31RBSFA | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | No voltage at bulb socket. Ran out of time to investigate further. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201525 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 25 | Microscope | Askmania/SLG Chemnitz | RME5 | 1193174 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Labelled specs for missing bulb, fixed hinge. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201526 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 26 | Scales (laboratory and in wards) | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Needed a cable tie to secure bassinet to scale. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201527 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 27 | Fetal steth (fetoscope or Doppler) | Biocare | FM-200 | 3314040595 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Broken speaker but working with audio jack output. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201528 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 28 | Exercise bicycle | Exercycle | H8705M | S6117539 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Repaired loose cables in original power jack. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201529 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 29 | Physiotherapy electrodes | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Repaired loose connections between electrode pads and cables with epoxy. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201530 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 30 | Spectrophotometer/Colorimeter | 2MM | Diagnostiko-550 | 4505-1001 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Needed a new bulb (quite complicated to replace). Printer had been turned off. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201531 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 31 | Microscope | LW Scientific | Revelation III | 128465 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Slide-holding mechanism was off its rails and bent out of shape. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201532 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 32 | Patient monitor | Inohero | H8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Poor connection on pulse oximeter socket. Didn't find time to return to work on it. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201533 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 33 | Patient monitor | Goldway | UT4000F | CN4FAAJW00562 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Bad pulse oximeter output from socket. Poor ecg signal. Didn't find time to return to work on it. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201534 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 34 | Patient monitor | Goldway | UT4000F | CN4FABAQ00624 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Bad pulse oximeter output from socket. Circuits and connections are all good, perhaps chips damaged by static electricity. Referred hospital to local servicer before leaving. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201535 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 35 | Ventilator | Neumovent | graphnet advance | 120710341081A1V | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Leak in filter for pressurised air input. Didn't find time to return to work on it. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201536 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 36 | Centrifuge | Clay Adams | Compact II | 14-04-4-59-3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Lid broken off, rejoined with epoxy. Seal missing, replaced with IV tubing. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201537 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 37 | Shaker Machine (laboratory) | Chem index inc | PS600T | 831260 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Pipette holder disconnected from shaking rod (missing screw), rejoined with with a ring of rubber tubing. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201538 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 38 | Oven (laboratory, not kitchen) | Rajas Scientific | Dry Heat Sterilizer | None | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 220V UK plug. Missing manual and unknown how to change desired temperature. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201539 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 39 | Nebulizer | Wellpro | NBA-02WD | 356 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | Fried inside. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201540 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 40 | Nebulizer | Wellpro | NBA-02WD | 355 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Working fine. Advised that manual says to give it half hour breaks after half an hour of use. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201541 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 41 | Electrosurgery Unit (ESU) | Birtcher | Hyfrecator 732 | B27443 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Only partially works/some of the time. Loose connection inside the pen. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201542 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 42 | Oven (laboratory, not kitchen) | Memmert | U10? | 14-04-4-101-1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Were told it wasn't regulating the temperature properly and that it burnt paper covered equipment. Seems fine but temperature dial needed to be recalibrated and digital thermometer readings are incorrect. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201543 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 43 | Nebulizer | DeVILBISS | PulmoAide 5650D | D4092803 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Working fine. Made a hole so case lid can be closed will cord is attached to not block air filter. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201544 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 44 | Nebulizer | Drive Medical | Spider Neb II | M70308730 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Piston disconnected from motor. Secured with cable tie to prevent coming off in the future. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201545 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 45 | Nebulizer | John Bunn | Neb-u-Lite | AJM296RP6160 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Piston broken, replaced with the one from Item 39. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201546 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 46 | Oven (laboratory, not kitchen) | FischerScientific | Isotemp 506G | 60800197 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Bothered that the numbers on the dial didn't represent the temperature settings. No changes made. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201547 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 47 | Rotator (laboratory) | Chem index inc | DSR 2100V | 831130153 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Broken belt, replaced with a ring of a blood pressure cuff. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201548 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 48 | Spectrophotometer/Colorimeter | Awareness Technology inc | Stat Fax 4542 | 4503-1134 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Wasn't printing the calculations they wanted, changed the settings with them. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201549 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 49 | Capillary centrifuge | Intramedica | Illegible | Illegible | 0 | 0 | 0 | 1 | 0 | 0 | 0 | #REF! | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201550 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 50 | Pipette | DragonLab | Multi 50-300uL | DX75176 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Rubber barrier preventing vials from hitting outer wall was old, replaced with a piece truck tyre tube. Spliced cable with one having three prongs. Replaced bypassed fuse. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201551 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 51 | Centrifuge | Chem index inc | DSC1512T | 2010310-22 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Working but stirring up rust particles. As much corrosion as possible covered with epoxy. Fuse changed to correct value (3A to 5A). | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201552 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 52 | Oven (laboratory, not kitchen) | Memmert | 530 | 660745 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Handle missing. Replaced with wire which hooked onto more wire on the side of the oven. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201553 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 53 | Oven (laboratory, not kitchen) | Lab-Line | 125 | 0198-0344 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Door clip damaged. Epoxied a hook and attached wire to the handle to act as a lock. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201554 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 54 | Spectrophotometer/Colorimeter | Rayto | RT-2100C | 451218015 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Wanted a printer (HP D2660) connected to it but it only accepts five particular models. | 0 | 1
Asunción, Juigalpa and surrounding health centres in Chontales20/07/201555 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 55 | ECG Machine | Cardioline | ar600adv | 05081123 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Somehow the language changed into english. The manual doesn't say how to change language and the manufacturer did not reply. Printed a translation sheet. | 1 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asunción, Juigalpa and surrounding health centres in Chontales20/07/20150 | Asunción, Juigalpa and surrounding health centres in Chontales | 20/07/2015 | Deanna Hood, Sindora Baddam | Nicaragua | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0