## Work Summary Form
|  |  | Engineering World Health | [8 empty] | 151 E. Rosemary St. Ste. 201
Chapel Hill, NC 27514
ewh.org
Complete electronically if possible
* REQUIRED FIELD
| *DATE | 20/07/2016 |  |  | *ENGINEER NAME(S) |  |  |  |  | Engineer Names 2
| *COUNTRY | Country 2 |  |  | *Hospital/CITY |  |  |  |  | Hospital 2
| EQUIPMENT INFORMATION |  |  |  | NATURE OF THE PROBLEM & SOLUTION (enter "1") | [7 empty] | Results |  | Date worked on / Date returned
| Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
GUIDE
1 | Aspirator/Suction Machine | Nakamura Medical Industry CO., LTD | GT-200 | 507006 |  |  |  | 1 |  |  |  | Missing tube connections, Obstruction found in the pressure sensor tube, Pressure sensor still not working, however machine is workable and was returned to service. | 1 |  | 21/06/17
2 | Nebulizer | DeVILBISS | 565OD | D4101915 | [6 empty] | 1 | Routine maintenance | 1 |  | 21/06/17
3 | Nebulizer | DeVILBISS | 565OD | D4101917 | [6 empty] | 1 | Routine maintenance | 1 |  | 21/06/17
4 | Nebulizer | MADA Medical | 180 | 0414NB14C00039 | [6 empty] | 1 | Routine maintenance | 1 |  | 22/06/17
5 | Nebulizer | DeVilbiss Healthcare | 5650C | 3C6001979 | [6 empty] | 1 | Routine maintenance | 1 |  | 22/06/17
6 | Ventilator | EKOM | DK50 DS | A1307-09-2015 | [5 empty] | 1 |  | Medical Air filter was connected poorly, tightened connection and instructed staff on how to connect properly to ensure optimum usage | 1 |  | 22/06/17
7 | Ventilator | TECME S.A. | 86 | 160740551081A1V | [5 empty] | 1 |  | Connected to above machine, medical air filter was improperly connected. Connection was fixed and tightened, then instructed staff on how to properly connect and operate. | 1 |  | 22/06/17
8 | Nebulizer | Sunrise | 4650D | D 2858029 | [6 empty] | 1 | Routine maintenance | 1 |  | 22/06/17
9 | Nebulizer | Sunrise | 3650D | D 7204359 | [6 empty] | 1 | Routine maintenance | 1 |  | 22/06/17
10 | Nebulizer | Sunrise | 5650D | D4126013 | [6 empty] | 1 | Routine maintenance | 1 |  | 22/06/17
11 | Phototherapy device | Fanem LTD. | 5006 | JAJ 10673 | [6 empty] | 1 | Routine maintenance | 1 |  | 23/06/17
12 | Incubator (infant) | Fanem LTD. | 1186 | CN6059 | [6 empty] | 1 | Routine maintenance | 1 |  | 23/06/17
13 | Centrifuge (electric or hand operated) | MG | 3411 | 804909 | [6 empty] | 1 | Routine maintenance | 1 |  | 23/06/17
14 | Other | Fuji Photo Film CO., LTD | CR-IR 362 | 56622036 2005 | [6 empty] | 1 | Xray film digitizer. Cleaning performed, two image slots are out of service and unrepairable without proper replacement part |  | 1 | 23/06/17
15 | Infant Warmer (radiant or other) | Fanem LTD. | 2051-UCI | YAF - 60133 | [6 empty] | 1 | Disassembly and thorough cleaning. Replaced lightbulb. | 1 |  | 26/06/17
16 | Infant Warmer (radiant or other) | Zhengzhou Dison Instrument and Meter Co., LTD. | BN-100 | 1131211015 | [6 empty] | 1 | Disassembly and thorough cleaning. Routine Maintenance. | 1 |  | 26/06/17
17 | Incubator (infant) | Fanem LTD. | 1186 | CN6066 | [6 empty] | 1 | Cleaning and routine maintenance. | 1 |  | 26/06/17
18 | ECG Machine | KONTRON Instruments | 7137B | 1933 |  |  |  | 1 |  |  |  | Clogged tubing. Cleaning performed and routine maintenance check. |  |  | 26/06/17
21 | Other | Chattanooga | 2778 | T10052 |  |  | 1 |  |  |  |  | Hospital Chicigalpa - Electric Muscle Stimulator Machine. Error with a computer chip canal. |  | 1 | 27/06/17
22 | Other | KaVo Dental Excellence |  | 2014115759 | 1 | [6 empty] | Hospital Chicigalpa - Dental Chair. Pressurized water not continually flowing into proper portals. Machine still works, needs further inspection. | 1 |  | 27/06/17
20 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI 5000 | 01ZHB8 |  |  |  | 1 |  |  |  | Hospital Corinto - Machine makes noise as temperature increases. Usable in air conditioning only. Cleaning performed, unable to remove necessary parts to try and fix the noise problem | 1 |  | 28/06/17
19 | Ultrasound Machine (imaging) | General Electric | 459 | 305486SU3 | [5 empty] | 1 |  | Ultrasound is functioning properly, thoroughly cleaned entire machine and ran maintanence tests. Connected image printer is not calibrated to machine properly, and doesn't print ultrasound images. | 1 |  | 29/06/17
23 | ECG Machine | KONTRON Instruments | 7137B | 1940 |  |  | 1 |  |  |  |  | Cleaned and unclogged tubes. Current flow exists, but digital screen is not functioning. |  | 1 | 29/06/17
24 | Drying Machine | [9 empty] | 1 | Routine maintenance - cleaning | 1 |  | 30/06/17
25 | Microscope | Olympus | CX31RBSF | 5AO3255 | [6 empty] | 1 | Routine maintenance - cleaning, replace lightbulb -- the connection to the light bulb was insecure, so a wire was added to improve the connection in case the lightbulb shifted | 1 |  | 2017-03-07
26 | Printer | HP | LaserJet P1102w | VNB3J90725 | [6 empty] | 1 | Routine maintenance - cleaning; toner damaged | 1 |  | 2017-03-07
27 | Patient Monitor | Philips Medical System | M8105A | DE21062420 | [6 empty] | 1 | Cleaning performed | 1 |  | 2017-05-07
28 | Patient Monitor | Philips Medical System | M8105A | DE91336900 | [6 empty] | 1 | Cleaning performed | 1 |  | 2017-05-07
29 | Patient Monitor | Goldway (US), Inc. | UT4000F | CN4FAAJR00050 | [6 empty] | 1 | Cleaning performed | 1 |  | 2017-05-07
30 | Pulse Oximeter | MasimoSet | RAD-5V | N22343 | [6 empty] | 1 | Replaced batteries and cleaned exterior of device | 1 |  | 2017-05-07
31 | Nebulizer | Sunrise Medical | 2655D | D1252720 | [6 empty] | 1 | Cleaning performed | 1 |  | 2017-05-07
32 | Nebulizer | WellPro | NBA-03WA | 311 | [6 empty] | 1 | Cleaning performed | 1 |  | 2017-05-07
33 | Microscope | Olympus Optical Co, LTD. | CH30RF100 | 8E07743 | [6 empty] | 1 | Cleaned and replaced lightbulb. | 1 |  | 2017-06-07
34 | Microscope | LW Scientific | Revelation III | 3336 | [6 empty] | 1 | Cleaned and replaced lightbulb | 1 |  | 2017-06-07
35 | Microscope | Westover Scientific | MC1200 | 11110021005 | [6 empty] | 1 | Cleaning and routine maintenance | 1 |  | 2017-06-07
36 | Centrifuge (electric or hand operated) | Miami Medical | 3411 | 908688 | [6 empty] | 1 | Tested machine operation. Timer dial is missing, but machine still functions correctly. Impossible to correctly set timer for desired time. | 1 |  | 2017-07-07
37 | Centrifuge (electric or hand operated) | Kokusan Corporation | H-1200F | 129177 | [5 empty] | 1 |  | Tested machine operation. Only operates with an attached voltage regulator. | 1 |  | 2017-07-07
38 | Other | Leisegang Medical, Inc. | GUL | 31924
39 | ECG Machine | HP | M1772A | CNA4002964 |  |  |  | 1 |  |  |  | The circuit board was dirt, and was interfering with the system's functionality | 1 |  | 2017-07-12
40 | Aspirator/Suction Machine | Dia-Pump | EF | 6256 | [6 empty] | 1 | Cleaning and routine maintenance | 1 |  | 2017-07-12
41 | Nebulizer | Mada Medical | 277 | [7 empty] | 1 | Cleaning and routine maintenance | 1 |  | 2017-07-12
42 | Aspirator/Suction Machine | Allied Healthcare Products, Inc. GOMCO | 6036 | P-6076 |  |  |  | 1 |  |  |  | Pressure sensor does not work, clamped off tube and cleaned machine. | 1 |  | 2017-07-14
43 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-3500 | 01QY83 | [6 empty] | 1 | Machine is in good working condition except for monitor. Image created is fuzzy and unreadable, a glitch in the graphics or software. Replaced monitor and cleaned up computer card. | 1 |  | 2017-07-20
44 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-3500 | 01G3JY | [5 empty] | 1 |  | Older of the two ultrasounds, took this monitor to replace above machine. This ultrasound works, but now has the faulty monitor. |  | 1 | 2017-07-20
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96
97
98
99
100

## Equipment Categories
Air Compressor
Air Conditioner
Anesthesia Machine
Aspirator/Suction Machine
Autoclave (lab, surgery, and other)
Automatic Voltage Regulator
Bed, delivery
Blood Bank Refrigerator
Blood clotting time meter
Blood electrolyte analyzer
Blood Gas Analyzer
Blood Pressure Device, Automatic (NIBP)
Blood Pressure Device, Manual (Sphygmomanometer) (Fixed or portable)
Bottle Washing Machine
Breast Pump
Capnograph
Ceiling Fan
Centrifuge
Centrifuge (electric or hand operated)
Change Over Switch
Computer
Control Switch Panel
Defibrillator (automatic and manual)
Dental Drilling Machine
Dialysis Equipment
Distiller
Drying Machine
ECG Machine
Electrosurgery Unit (ESU)
Fetal steth (fetoscope or Doppler)
Fluoroscopy (x-ray moving images)
Furniture (chairs, tables, and beds)
Generator
Glucose level kit (or glucometer)
Heart Lung Machine
Hot Plate (laboratory, but not kitchen)
Incubator (infant)
Infant Warmer (radiant or other)
Infusion pumps
Iron (for clothing)
Lamp, examination
Lamp, surgical
Laryngoscope
Microscope
Nebulizer
Operating Table
Ophthalmoscope (eye examination instrument)
Otoscopes
Oven (laboratory, not kitchen)
Oxygen Concentrator
Pacemaker (ext, temp and permanent)
Patient Monitor
Photocopier
Phototherapy device
Printer
Projector
Pulse Oximeter
Respiration monitor (apnea monitor)
Scales (laboratory and in wards)
Shaker Machine (laboratory)
Skin Grafting Machine (dermatome)
Spectrophotometer/Colorimeter
Stethoscopes
Telephone
Television
Thermometers
Transformer
Ultrasound Machine (imaging)
UPS (battery backup for computer)
Vaccine Refrigerator
Vacuum Extractor (for delivery)
Ventilator
Washing Machine
Water Bath (laboratory)
Water Pump (for drinking water)
Water Purifier (for lab, in wards)
X-Ray Film Dryer
X-Ray Film View Box
X-Ray Machine
Other

## Reports
Repair ID | Hospital | Date | Engineers | Country | Entry # | Equipment TYPE (select the type from the EWH equipment types) | Manufacturer | Model | Serial Number | Plumbing | Motor | Electric Simple | Mechanical Simple | Power Supply | Isntallation/Training | Other | Notes | Repaired | Abandoned
Hospital 220/07/20161 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 1 | Aspirator/Suction Machine | Nakamura Medical Industry CO., LTD | GT-200 | 507006 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Missing tube connections, Obstruction found in the pressure sensor tube, Pressure sensor still not working, however machine is workable and was returned to service. | 1 | 0
Hospital 220/07/20162 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 2 | Nebulizer | DeVILBISS | 565OD | D4101915 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/20163 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 3 | Nebulizer | DeVILBISS | 565OD | D4101917 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/20164 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 4 | Nebulizer | MADA Medical | 180 | 0414NB14C00039 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/20165 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 5 | Nebulizer | DeVilbiss Healthcare | 5650C | 3C6001979 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/20166 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 6 | Ventilator | EKOM | DK50 DS | A1307-09-2015 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Medical Air filter was connected poorly, tightened connection and instructed staff on how to connect properly to ensure optimum usage | 1 | 0
Hospital 220/07/20167 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 7 | Ventilator | TECME S.A. | 86 | 160740551081A1V | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Connected to above machine, medical air filter was improperly connected. Connection was fixed and tightened, then instructed staff on how to properly connect and operate. | 1 | 0
Hospital 220/07/20168 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 8 | Nebulizer | Sunrise | 4650D | D 2858029 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/20169 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 9 | Nebulizer | Sunrise | 3650D | D 7204359 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/201610 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 10 | Nebulizer | Sunrise | 5650D | D4126013 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/201611 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 11 | Phototherapy device | Fanem LTD. | 5006 | JAJ 10673 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/201612 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 12 | Incubator (infant) | Fanem LTD. | 1186 | CN6059 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/201613 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 13 | Centrifuge (electric or hand operated) | MG | 3411 | 804909 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance | 1 | 0
Hospital 220/07/201614 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 14 | Other | Fuji Photo Film CO., LTD | CR-IR 362 | 56622036 2005 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Xray film digitizer. Cleaning performed, two image slots are out of service and unrepairable without proper replacement part | 0 | 1
Hospital 220/07/201615 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 15 | Infant Warmer (radiant or other) | Fanem LTD. | 2051-UCI | YAF - 60133 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Disassembly and thorough cleaning. Replaced lightbulb. | 1 | 0
Hospital 220/07/201616 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 16 | Infant Warmer (radiant or other) | Zhengzhou Dison Instrument and Meter Co., LTD. | BN-100 | 1131211015 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Disassembly and thorough cleaning. Routine Maintenance. | 1 | 0
Hospital 220/07/201617 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 17 | Incubator (infant) | Fanem LTD. | 1186 | CN6066 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning and routine maintenance. | 1 | 0
Hospital 220/07/201618 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 18 | ECG Machine | KONTRON Instruments | 7137B | 1933 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Clogged tubing. Cleaning performed and routine maintenance check. | 0 | 0
Hospital 220/07/201619 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 19 | Ultrasound Machine (imaging) | General Electric | 459 | 305486SU3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Ultrasound is functioning properly, thoroughly cleaned entire machine and ran maintanence tests. Connected image printer is not calibrated to machine properly, and doesn't print ultrasound images. | 1 | 0
Hospital 220/07/201620 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 20 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI 5000 | 01ZHB8 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Hospital Corinto - Machine makes noise as temperature increases. Usable in air conditioning only. Cleaning performed, unable to remove necessary parts to try and fix the noise problem | 1 | 0
Hospital 220/07/201621 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 21 | Other | Chattanooga | 2778 | T10052 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Hospital Chicigalpa - Electric Muscle Stimulator Machine. Error with a computer chip canal. | 0 | 1
Hospital 220/07/201622 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 22 | Other | KaVo Dental Excellence | 0 | 2014115759 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | Hospital Chicigalpa - Dental Chair. Pressurized water not continually flowing into proper portals. Machine still works, needs further inspection. | 1 | 0
Hospital 220/07/201623 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 23 | ECG Machine | KONTRON Instruments | 7137B | 1940 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | Cleaned and unclogged tubes. Current flow exists, but digital screen is not functioning. | 0 | 1
Hospital 220/07/201624 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 24 | Drying Machine | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance - cleaning | 1 | 0
Hospital 220/07/201625 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 25 | Microscope | Olympus | CX31RBSF | 5AO3255 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance - cleaning, replace lightbulb -- the connection to the light bulb was insecure, so a wire was added to improve the connection in case the lightbulb shifted | 1 | 0
Hospital 220/07/201626 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 26 | Printer | HP | LaserJet P1102w | VNB3J90725 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Routine maintenance - cleaning; toner damaged | 1 | 0
Hospital 220/07/201627 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 27 | Patient Monitor | Philips Medical System | M8105A | DE21062420 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning performed | 1 | 0
Hospital 220/07/201628 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 28 | Patient Monitor | Philips Medical System | M8105A | DE91336900 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning performed | 1 | 0
Hospital 220/07/201629 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 29 | Patient Monitor | Goldway (US), Inc. | UT4000F | CN4FAAJR00050 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning performed | 1 | 0
Hospital 220/07/201630 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 30 | Pulse Oximeter | MasimoSet | RAD-5V | N22343 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Replaced batteries and cleaned exterior of device | 1 | 0
Hospital 220/07/201631 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 31 | Nebulizer | Sunrise Medical | 2655D | D1252720 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning performed | 1 | 0
Hospital 220/07/201632 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 32 | Nebulizer | WellPro | NBA-03WA | 311 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning performed | 1 | 0
Hospital 220/07/201633 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 33 | Microscope | Olympus Optical Co, LTD. | CH30RF100 | 8E07743 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaned and replaced lightbulb. | 1 | 0
Hospital 220/07/201634 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 34 | Microscope | LW Scientific | Revelation III | 3336 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaned and replaced lightbulb | 1 | 0
Hospital 220/07/201635 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 35 | Microscope | Westover Scientific | MC1200 | 11110021005 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning and routine maintenance | 1 | 0
Hospital 220/07/201636 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 36 | Centrifuge (electric or hand operated) | Miami Medical | 3411 | 908688 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Tested machine operation. Timer dial is missing, but machine still functions correctly. Impossible to correctly set timer for desired time. | 1 | 0
Hospital 220/07/201637 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 37 | Centrifuge (electric or hand operated) | Kokusan Corporation | H-1200F | 129177 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Tested machine operation. Only operates with an attached voltage regulator. | 1 | 0
Hospital 220/07/201638 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 38 | ECG Machine | HP | M1772A | CNA4002964 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | The circuit board was dirt, and was interfering with the system's functionality | 1 | 0
Hospital 220/07/201639 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 39 | Aspirator/Suction Machine | Dia-Pump | EF | 6256 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning and routine maintenance | 1 | 0
Hospital 220/07/201640 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 40 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Hospital 220/07/201641 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 41 | Nebulizer | Mada Medical | 277 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Cleaning and routine maintenance | 1 | 0
Hospital 220/07/201642 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 42 | Aspirator/Suction Machine | Allied Healthcare Products, Inc. GOMCO | 6036 | P-6076 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | Pressure sensor does not work, clamped off tube and cleaned machine. | 1 | 0
Hospital 220/07/201643 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 43 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-3500 | 01QY83 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | Machine is in good working condition except for monitor. Image created is fuzzy and unreadable, a glitch in the graphics or software. Replaced monitor and cleaned up computer card. | 1 | 0
Hospital 220/07/201644 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 44 | Ultrasound Machine (imaging) | ATL Ultrasound | HDI-3500 | 01G3JY | 0 | 0 | 0 | 0 | 0 | 1 | 0 | Older of the two ultrasounds, took this monitor to replace above machine. This ultrasound works, but now has the faulty monitor. | 0 | 1
Hospital 220/07/201645 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 45 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201646 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 46 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201647 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 47 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201648 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201649 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 49 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201650 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201651 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201652 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 52 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201653 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201654 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 54 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201655 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 55 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201656 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 56 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201657 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 57 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201658 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 58 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201659 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 59 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201660 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 60 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201661 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 61 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201662 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 62 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201663 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201664 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 64 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201665 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201666 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 66 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201667 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201668 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201669 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 69 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201670 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 70 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201671 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 71 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201672 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 72 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201673 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 73 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201674 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201675 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201676 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 76 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201677 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201678 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 78 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201679 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 79 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201680 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 80 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201681 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 81 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201682 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 82 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201683 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 83 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201684 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 84 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201685 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 85 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201686 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 86 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201687 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 87 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201688 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 88 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201689 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201690 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201691 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201692 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 92 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201693 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 93 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201694 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 94 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201695 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 95 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201696 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 96 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201697 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 97 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201698 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 98 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/201699 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 99 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hospital 220/07/2016100 | Hospital 2 | 20/07/2016 | Engineer Names 2 | Country 2 | 100 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0